## Slide 1: Thin Film Microanalysis Using the Electron Microprobe
Paul Carpenter
Earth and Planetary Sciences CB 1169
Washington University 1 Brookings Drive
St. Louis, MO 63130
paulc@levee.wustl.edu

## Slide 2: EPMA and thin film analysis
Conventional EPMA: Characteristic x-rays measured on sample vs. standard, comparative method “k-ratio” : background-subtracted x-ray intensity from sample v. standard: k = (P-B)sample / (P-B)standard Measured k-ratios are converted to concentration units: Csample = Cstd K * Z A F Z = Atomic number correction (electron energy loss, backscattering) A = Absorption correction (x-ray absorption in sample and standard) F = Fluorescence (nominally characteristic, can include continuum)
Method assumes homogeneous bulk structure of sample and standard
Thin film algorithm is enhancement of F(rz) code Assumption is homogeneous bulk standard for convienience Sample multilayer structure – user provides model (e.g., Cu/Ni/Au layer) Algorithm treats electron behavior, and x-ray generation, absorption, and fluorescence in multilayer structure
Results provide chemical analysis and layer thickness Analytical summation reflects model validity and measurement quality

## Slide 3: Multi-kV BSE imaging: Stillwater Ca-pyroxene with low-Ca pyroxene exsolution lamellae, 350x mag

## Slide 4: Thin-film analysis
TF: multilayer stratified materials analyzed in plan view
X-ray fluorescence spectrometry : bulk, high sensitivity
EPMA: multi-kV and F(rz) algorithm coupled with model High spatial resolution Very thin film (monolayer) measurement possible Both EDS and WDS can be used
Multilayer materials can be modeled using Monte Carlo simulation Casino – excellent visualization DTSA-II – MC, F(rz), geometries, scattering, EDS, char+cont fluor. Penelope – Advanced software, complete simulation, cpu intensive
EPMA F(rz) technique calibrated using multilayer materials
TF software using F(rz) algorithm for multilayer structure: Stratagem: Pouchou and Pichoir algorithm, Windows, ~5k GMRfilm: PAP algorithm, PCDOS, free Oxford ThinFilmID, EDS, Inca
Multilayer

## Slide 5: Casino Monte Carlo Program – Bulk Sample
Simulation Fe3C, 15 keV, 5000 trajectories, max depth ~750nm Red = backscattered e-, blue e- within Fe3C. Scattering angle increases as e- energy decreases (curly trajectory)

## Slide 6: Casino MC Simulation of F(rz) Experiment
Ni matrix
Ni overlayer 25nm
Cu layer 25nm
F(rz)
Depth rz
Generated F(rz) for Cu layer
Generated F(rz) for Ni overlayer and matrix
The F(rz) distribution for Cu Ka is obtained from the emitted intensity of a layer of Cu sandwiched in a matrix with similar Z (Ni) so that scattering properties are similar. Series of measurements are made for progressively buried layers. Cu intensity is divided by that from an isolated Cu foil to yield the F(rz) curve. The curves are then generalized to model x-ray behavior for elements in general.

## Slide 7: Casino Monte Carlo Program – Multilayer Sample
Sample 250 nm Aluminum layer on Fe substrate Simulation at 15 keV, 25000 trajectories, max depth ~920nm Red = backscattered e-

## Slide 8: Thin film samples: layer/film on metal substrate EDS spectra at multiple kV
5 kV
Oxide layer formed on polished Mg metal Oxygen Ka peak intensity increases with decrease in accelerating voltage
C
O
Mg
20 kV
Solvent residue on polished brass Carbon and Oxygen present Layer not obvious at normal operating conditions of 15-20 kV
C
O
Cu + Zn L-lines
Cu + Zn K-lines

## Slide 9: Casino – Multilayer Sample 250 nm Al on Fe
Al Ka F(rz) multilayer
Fe Ka F(rz) multilayer
Al Ka F(rz) bulk Al
Fe Ka F(rz) bulk

## Slide 10: GMRfilm output: 250 nm Al on Fe substrate
GMRfilm output for 2500 Angstrom Al layer on Fe substrate
Here the structure is described and expected k-ratio is calculated for Al and Fe
That is, from the Al film we should observe 20% emitted intensity relative to pure Al standard
Run the program for multiple kV to get data for comparison with AlFe bulk sample

## Slide 11: TF Al on Fe vs. Al50-Fe50 bulk sample
GMRfilm k-ratios for 2500 Angstrom Al layer on Fe substrate, compared with CalcZAF k-ratios for Al-Fe bulk sample (Al 32.6, Fe 67.4 wt%)
Stratified sample: k-ratio function of kV, some low kV where topmost layer is sampled
Bulk sample k-ratio is not strong function of kV
Greatest sensitivity of k-ratio to film thickness is observed when scattering volume transits the layer structure (5-15kV). 20-25 kV data not good for accurate film measurement of Al on Fe
This modeling procedure also used by Stratagem—best results from multiple kV data
In practice, TF programs calculate thickness and composition from single or multiple kV measurements

## Slide 12: DTSA-II Program
Simulation of EDS spectra and scattering volume using F(rz) and Monte Carlo codes
Geometries: Bulk, Multilayer thin film, Sphere/cube on substrate, Cube embedded in substrate, Interface (grain boundary)
Bulk
Interface
Sphere on substrate
Cube on substrate
Cube in substrate
Multilayer

## Slide 13: EDS spectra: ~ 1 mm ZnTe film on GaAs
ZnTe on GaAs wafer, 1 inch round
Ga and As L-lines not observed at any kV Ga La macs by ZnTe 8006 As La macs by ZnTe 5785 ZnTe overlayer essentially opaque to Ga,As L-lines
Ga and As K-lines observed above ~ 20 kV As Ka absorbed by Ga (mac ~200) But K-lines penetrate easily through ZnTe overlayer
Ga
As
Zn
Zn

## Slide 14: DTSA-II Simulation: ZnTe film on GaAs substrate
Simulation consistent with ~ 1 mm film thickness
Zn
Ga
As
Simulation of film thickness 0.1 – 2 mm
20 kV
Experimental EDS spectra do not show predicted Ga and As L-lines
K Lines
L Lines

## Slide 15: Pt 1 mm thin film on Ni substrate
Pt
Ni
Primary E0 > 30 kV
Secondary fluorescence E0 > 10 kV
Excitation of Ni Ka in substrate due to: 1. Primary beam penetration through Pt film (> 30kV) 2. Characteristic fluorescence from Pt La (> ~12kV) 3. Continuum fluorescence (>Ec Ni Ka) Pouchou and Pichoir, 1991 Electron Probe Quantitation
Casino simulation 20 kV (Casino v2.48) Characteristic and continuum fluorescence not treated No Ni production predicted below ~30 kV, incorrect

## Slide 16: DTSA-II Simulation: Pt 1 um film on Ni substrate
Excitation of Ni Ka in TF example is entirely due to secondary fluorescence below 30 kV No Ni La observed due to high absorption in Pt overlayer
30 kV 25 20 15
Not including characteristic and continuum fluorescence
Including characteristic and continuum fluorescence

## Slide 17: DTSA-II Scattering diagrams Pt 1 mm on Ni @ 20 kV
Electron scattering plot
Pt La excitation in layer by electrons only
Ni Ka excitation in substrate by Pt La and continuum x-rays

## Slide 18: Conclusions
EPMA/SEM thin film analysis Excellent vertical sensitivity Thickness and composition determined Can in principle detect very thin layers
Excellent software tools for visualization of scattering volume Simulation with and without x-ray fluorescence
For materials of unknown layer structure Methods for deducing multilayer structure using these tools
Role of x-ray fluorescence important Continuum and characteristic (Pt film on Ni example)

## Slide 19: DTSA-II Simulation: ZnTe film on GaAs substrate