## Data
Detector Bias and PHA Data Excel Worksheet
Paul Carpenter
Washington University
Earth and Planetary Sciences, Campus Box 1169, 1 Brookings Drive, St. Louis, MO 63130
paulc@levee.wustl.edu
NOTE: May need to use count rate of 20K cps or shifts are observed (especially for H-type spec using PET crystal for Ti at 20KV)
For bias scans it is necessary to scan from -50v to +50v (i.e. 1550-1650 for nominal 1600v) relative to these nominal bias values at a given gain
(Using JEOL software, suggest 2-5v step size and 200 msec for quick and 500 msec for precision scans)
Using PFW, Startwin for bias scans all spectrometers, use 1 second count time, typical bias range 1500-1850 general
Detector Bias and PHA Data Excel sheet from a JEOL 8530 lab
Measured values in bold, other values adjusted to conform to linear fit are in italics
Spectrometer configuration
| Spec 1 | Spec 2 | Spec 3 | Spec 4 | Spec 5
| TAPH | PETL | LDE1L | PETL | LiFH |  | To update the charts, enter new bias values
| LDE6H | TAPL | TAPL | LiFL | PETH |  | here in appropriate column
LDE6H for B C N | [6 empty] | LS fit will be updated, charts updated
LDE1L for O F | [6 empty] | And line plotted on charts from next worksheet |  |  |  |  | NOTE:
The slope and intercept values are copied to
| XPCH | XPCH | GPCH | GPCH | GPCH |  | To edit the linest function, select both cells |  |  |  |  | next worksheet for table calculation
[7 empty] | Edit in formula bar and use ctrl+shift+enter |  |  |  |  | Least Square Fit to Bias Data
Spectrometer 1 | X-ray Line | Peak pos | Energy Kev | ln Energy | Std | 8x Bias | 16x Bias | 32x Bias | 64x Bias | 128x Bias |  |  | Slope | Intercept
LDE6H | B Ka | 158.62 | 0.183 | -1.6983 |  |  |  |  | 1735 | 1660 |  | 8x | -79.3133 | 1789.415
| C Ka | 104.79 | 0.277 | -1.2837 |  |  |  | 1760 | 1685 | 1615 |  | 16x | -80.8085 | 1726.6126
TAPH | Fe La | 191.27 | 0.705 | -0.34956 |  |  | 1755 | 1709 | 1637 | 1574 |  | 32x | -75.5327 | 1668.8261
| Cu La | 144.97 | 0.93 | -0.072571 |  |  | 1731 | 1673 | 1606 | 1544 |  | 64x | -79.3705 | 1595.6117
| Ga La | 122.753 | 1.098 | 0.09349 |  | 1782 | 1721 | 1662 | 1585 | 1523 |  | 128x | -68.983 | 1537.5244
| Al Ka | 90.68 | 1.486 | 0.39609 |  | 1758 | 1694 | 1632 | 1557
Two points for LS lines
[7 empty] | 16x | 32x | 64x | 128x
|  |  | 0.1 | -2.3026 |  |  | 1912.6811 | 1842.7465 | 1778.369 | 1696.3636
|  |  | 15 | 2.7081 |  |  | 1507.7791 | 1464.2798 | 1380.6724 | 1350.715
Spectrometer 2 | X-ray Line | Peak pos | Energy Kev | ln Energy |  | 8x Bias | 16x Bias | 32x Bias | 64x Bias | 128x Bias |  |  | Slope | Intercept
TAPL | Fe La | 191.27 | 0.705 | -0.34956 |  |  | 1749 | 1690 | 1630 | 1565 |  | 8x | -93.6514 | 1790.5934
| Cu La | 144.97 | 0.93 | -0.072571 |  |  | 1729 | 1670 | 1602 | 1538 |  | 16x | -88.72 | 1721.2155
| Ga La | 122.753 | 1.098 | 0.09349 |  | 1777 | 1714 | 1655 | 1584 | 1523 |  | 32x | -93.0116 | 1661.034
| Al Ka | 90.68 | 1.486 | 0.39609 |  | 1756 | 1688 | 1624 | 1554 |  |  | 64x | -104.4082 | 1593.9625
PETL | Si Ka | 228.22 | 1.739 | 0.55331 |  | 1741 | 1672 | 1609 | 1535 |  |  | 128x | -95.0718 | 1531.5853
| Ca Ka | 107.57 | 3.69 | 1.3056 |  | 1678 | 1607 | 1540
| Ti Ka | 87.95 | 4.508 | 1.5059 |  | 1640 | 1585 | 1520
Two points for LS lines
[7 empty] | 16x | 32x | 64x | 128x
|  |  | 0.1 | -2.3026 |  |  | 1925.501 | 1875.2011 | 1834.3713 | 1750.4962
|  |  | 15 | 2.7081 |  |  | 1480.9572 | 1409.154 | 1311.2198 | 1274.1261
Spectrometer 3 | X-ray Line | Peak pos | Energy Kev | ln Energy |  | 8x Bias | 16x Bias | 32x Bias | 64x Bias | 128x Bias |  |  | Slope | Intercept
LDE1L | C ka | 208.88 | 0.277 | -1.2837 |  |  |  | 1760 | 1710 | 1640 |  | 8x | -72.7038 | 1788.7971
| O ka | 110.2 | 0.525 | -0.64436 |  |  |  | 1715 | 1650 | 1590 |  | 16x | -75.7145 | 1722.7767
TAPL | Fe La | 191.27 | 0.705 | -0.34956 |  |  | 1747 | 1694 | 1629 | 1566 |  | 32x | -78.1205 | 1662.7743
| Cu La | 144.97 | 0.93 | -0.072571 |  |  | 1731 | 1668 | 1604 | 1538 |  | 64x | -92.6594 | 1592.4323
| Ga La | 122.753 | 1.098 | 0.09349 |  | 1782 | 1717 | 1657 | 1581 | 1517 |  | 128x | -87.6789 | 1530.6264
| Al Ka | 90.68 | 1.486 | 0.39609 |  | 1760 | 1691 | 1628 | 1553
Two points for LS lines
[7 empty] | 16x | 32x | 64x | 128x
|  |  | 0.1 | -2.3026 |  |  | 1897.1159 | 1842.6533 | 1805.7884 | 1732.5146
|  |  | 15 | 2.7081 |  |  | 1517.738 | 1451.22 | 1341.5061 | 1293.1875
Spectrometer 4 | X-ray Line | Peak pos | Energy Kev | ln Energy |  | 8x Bias | 16x Bias | 32x Bias | 64x Bias | 128x Bias |  |  | Slope | Intercept
PETL | Si Ka | 228.22 | 1.739 | 0.55331 |  |  |  | 1755 | 1694 | 1622 |  | 8x | #VALUE! | #VALUE!
| Ca Ka | 107.57 | 3.69 | 1.3056 |  |  | 1770 | 1693 | 1608 | 1535 |  | 16x | -105.3041 | 1909.2458
LiFL | Ti Ka | 191.166 | 4.508 | 1.5059 |  |  | 1750 | 1665 | 1585 |  |  | 32x | -96.7601 | 1812.5124
| Fe Ka | 134.715 | 6.4 | 1.8563 |  |  | 1719 | 1637 | 1552 |  |  | 64x | -110.5487 | 1753.9232
| Cu Ka | 115.28 | 8.04 | 2.0844 |  |  | 1691 | 1612 | 1523 |  |  | 128x | -115.6429 | 1685.9864
| Ga Ka | 93.27 | 9.244 | 2.224 |  | 1764 | 1671 | 1591
Note for Spec 4 PET could probably use 8x gain
Two points for LS lines
[7 empty] | 16x | 32x | 64x | 128x
|  |  | 0.1 | -2.3026 |  |  | 2151.7175 | 2035.3107 | 2008.471 | 1952.2639
|  |  | 15 | 2.7081 |  |  | 1624.0769 | 1550.4813 | 1454.5516 | 1372.8197
Spectrometer 5 | X-ray Line | Peak pos | Energy Kev | ln Energy |  | 8x Bias | 16x Bias | 32x Bias | 64x Bias | 128x Bias |  |  | Slope | Intercept
PETH | Si Ka | 228.22 | 1.739 | 0.55331 |  |  |  | 1792 | 1705 | 1631 |  | 8x | #VALUE! | #VALUE!
| Ca Ka | 107.57 | 3.69 | 1.3056 |  |  | 1782 | 1705 | 1621 | 1545 |  | 16x | -105.8377 | 1923.4037
LiFH | Ti Ka | 191.166 | 4.508 | 1.5059 |  |  | 1765 | 1685 | 1602 | 1525 |  | 32x | -111.0632 | 1852.5627
| Fe Ka | 134.715 | 6.4 | 1.8563 |  |  | 1733 | 1649 | 1561 |  |  | 64x | -111.1458 | 1766.9955
| Cu Ka | 115.28 | 8.04 | 2.0844 |  |  | 1703 | 1622 | 1534 |  |  | 128x | -112.1133 | 1692.7459
| Ga Ka | 93.27 | 9.244 | 2.224 |  | 1778 | 1684 | 1604
Two points for LS lines
[7 empty] | 16x | 32x | 64x | 128x
|  |  | 0.1 | -2.3026 |  |  | 2167.104 | 2108.2951 | 2022.9181 | 1950.8963
|  |  | 15 | 2.7081 |  |  | 1636.7899 | 1551.7979 | 1466.0071 | 1389.1375
Escape peaks observed when E line > Ec of counter gas, generate ionized gas atom, Ar Ka or Xe La photon emitted, lost voltage equivalent to that x-ray
| Gas | E line | Ec
| Ar | 2.957 | 3.203
| Xe | 4.111 | 5.452
Generation of Xe La escape peak, observe Xe La escape peak only for Eline > 5.452, i.e. K lines of Mn and higher, L line of Sm and higher, No M lines
Observed voltage is V = K * (E - Ect) where K is amp scaling value, E is line pulse voltage, and Ect is escape peak value
Using 1700 volt bias settings for all, if not K varies due to bias
| X-ray Line | Peak pos | Energy Kev | ln Energy |  |  |  | U rel to Xe |  | LIF4 esc, V | K
| Sb La | 239.16 | 3.605 | 1.2823 |  |  |  | 0.66123
| Ti Ka | 191.12 | 4.51 | 1.5063 |  |  |  | 0.82722
| Fe Ka | 134.62 | 6.403 | 1.8568 |  |  |  | 1.1744 |  | 2.25 | 1.0187
| Ni Ka | 115.28 | 7.477 | 2.0118 |  |  |  | 1.3714 |  | 3.2 | 1.0519
| Ga Ka | 93.18 | 9.251 | 2.2247 |  |  |  | 1.6968 |  | 5.6 | 0.91786

## Table
Calculated Bias Values as a Function of L Position, mm, for Spectrometers and Gain Values
Note: Use same settings for LIF and PET
n*lamda = 2dsin(theta), L = R/d * n * lamda (R is 140 mm for std Jeol, 100 mm for H spectrometer)
| Lamda to E | 12.3985 | Lamda = 12.39854141/E
| R | 140
| TAP 2d | 25.757 | Keep bold data, used in formulas
| PET 2d | 8.742
| LIF 2d | 4.027
NOTE: This worksheet uses slope and intercept values from LS fit previous worksheet page
|  |  | TAP Spectrometer Data | [14 empty] | PET Spectrometer Data | [20 empty] | LIF Spectrometer Data
Color rationale: Green for bias values in range 1550-1750
|  |  |  | Spectrometer 1 |  |  |  | Spectrometer 2 |  |  |  | Spectrometer 3 | [6 empty] | Spectrometer 2 |  |  |  |  | Spectrometer 4 |  |  |  | Spectrometer 5 | [11 empty] | Spectrometer 4 |  |  | Spectrometer 5 |  |  |  | None |  |  |  | None
Symbol | TAP, mm | Energy | ln(E) | 16x | 32x | 64x | 128x | 8x | 16x | 32x | 64x | 16x | 32x | 64x | Symbol | PET, mm | Energy | ln(E) | 8x | 16x | 32x | 64x | 128x | 16x | 32x | 64x | 128x | 16x | 32x | 64x | 128x | 16x | 32x | 64x | 128x | Symbol | LIF, mm | Energy | ln(E) | 16x | 32x | 64x | 16x | 32x | 64x | 128x | 16x | 32x | 64x | 128x | 16x | 32x | 64x | 128x
Be | [14 empty] | Be | [20 empty] | Be
B | LDE | 0.183 | -1.6983 | 1863.8473 | 1797.1009 | 1730.4041 | 1654.676 | 1949.6386 | 1871.886 | 1818.9927 | 1771.2757 | 1851.3604 | 1795.4439 | 1749.7928 | B | [20 empty] | B
C | LDE | 0.277 | -1.2837 | 1830.3496 | 1765.7903 | 1697.5026 | 1626.0804 | 1910.8172 | 1835.1088 | 1780.4365 | 1727.9953 | 1819.9743 | 1763.0605 | 1711.3826 | C | [20 empty] | C
N | :LDE | 0.392 | -0.93649 | 1802.2893 | 1739.562 | 1669.9416 | 1602.1265 | 1878.2973 | 1804.3012 | 1748.1387 | 1691.7401 | 1793.6829 | 1735.9336 | 1679.2072 | N | [20 empty] | N
O | 256.8481 | 0.52476 | -0.64482 | 1778.7198 | 1717.5313 | 1646.7916 | 1582.0062 | 1850.982 | 1778.4242 | 1721.01 | 1661.2873 | 1771.5992 | 1713.1481 | 1652.1812 | O | [20 empty] | O
F | 199.1806 | 0.67668 | -0.39055 | 1758.1724 | 1698.3254 | 1626.6098 | 1564.4657 | 1827.1689 | 1755.8651 | 1697.3597 | 1634.7391 | 1752.347 | 1693.2842 | 1628.6204 | F | [20 empty] | F
Ne | 158.8284 | 0.8486 | -0.16416 | 1739.8783 | 1681.2257 | 1608.6413 | 1548.8488 | 1805.9674 | 1735.78 | 1676.303 | 1611.1024 | 1735.2062 | 1675.5987 | 1607.6435 | Ne | [20 empty] | Ne
Na | 129.5344 | 1.0405 | 0.039716 | 1723.4033 | 1665.8263 | 1592.4594 | 1534.7847 | 1786.8739 | 1717.6919 | 1657.34 | 1589.8158 | 1719.7697 | 1659.6716 | 1588.7523 | Na | [20 empty] | Na
Mg | 107.5321 | 1.2534 | 0.22587 | 1708.3602 | 1651.7654 | 1577.6841 | 1521.943 | 1769.4401 | 1701.1761 | 1640.0253 | 1570.3796 | 1705.6749 | 1645.129 | 1571.5031 | Mg | [20 empty] | Mg
Al | 90.6828 | 1.4863 | 0.39629 | 1694.5886 | 1638.8929 | 1564.1576 | 1510.1868 | 1753.4798 | 1686.0562 | 1624.174 | 1552.5861 | 1692.7715 | 1631.8155 | 1555.7119 | Al | [20 empty] | Al
Si | 77.4973 | 1.7392 | 0.55342 | 1681.8917 | 1627.0249 | 1551.6866 | 1499.3479 | 1738.7649 | 1672.1161 | 1609.5596 | 1536.181 | 1680.8749 | 1619.5409 | 1541.1529 | Si | 228.0192 | 1.7416 | 0.5548 | 1738.6356 | 1671.9936 | 1609.4312 |  |  | 1850.8231 | 1758.8299 | 1692.5907 |  | 1864.685 | 1790.9448 | 1705.3318 | [5 empty] | Si
P | 66.954 | 2.0131 | 0.69966 | 1670.0744 | 1615.9791 | 1540.0796 | 1489.26 | 1725.0696 | 1659.1419 | 1595.9578 | 1520.9126 | 1669.8026 | 1608.1167 | 1527.6026 | P | 196.9977 | 2.0158 | 0.70104 | 1724.9403 | 1659.0194 | 1595.8294 |  |  | 1835.4237 | 1744.6799 | 1676.4243 |  | 1849.2075 | 1774.7032 | 1689.0781 | [5 empty] | P
S | [14 empty] | S | 171.9035 | 2.3101 | 0.8373 | 1712.1795 | 1646.9306 | 1583.1558 |  |  | 1821.0751 | 1731.4956 | 1661.3611 |  | 1834.7862 | 1759.5699 | 1673.9336 | [5 empty] | S
Cl | [14 empty] | Cl | 151.3171 | 2.6244 | 0.96485 | 1700.2337 | 1635.6138 | 1571.2916 |  |  | 1807.6429 | 1719.1532 | 1647.26 |  | 1821.286 | 1745.4031 | 1659.7562 | [5 empty] | Cl
Ar | [14 empty] | Ar | 134.1289 | 2.9607 | 1.0854 | 1688.9416 | 1624.9163 | 1560.0766 |  |  | 1794.9457 | 1707.4863 | 1633.9304 |  | 1808.5245 | 1732.0115 | 1646.3547 | [5 empty] | Ar
K | [14 empty] | K | 119.7204 | 3.317 | 1.1991 | 1678.2988 | 1614.8339 | 1549.5066 |  |  | 1782.9787 | 1696.4902 | 1621.3674 |  | 1796.4968 | 1719.39 | 1633.7238 | [5 empty] | K | 260.1518 | 3.3138 | 1.1981 | 1783.0829 | 1696.5859 | 1621.4768 | 1796.6015 | 1719.4999 | 1633.8338
Ca | [14 empty] | Ca | 107.463 | 3.6954 | 1.3071 | 1668.1833 | 1605.251 | 1539.4602 |  |  | 1771.6046 | 1686.0389 | 1609.4268 |  | 1785.065 | 1707.3938 | 1621.7187 | [5 empty] | Ca | 233.5165 | 3.6917 | 1.3061 | 1771.7087 | 1686.1347 | 1609.5361 | 1785.1697 | 1707.5037 | 1621.8287
Sc | [14 empty] | Sc | 96.982 | 4.0947 | 1.4097 | 1658.5726 | 1596.1465 | 1529.9152 |  |  | 1760.7981 | 1676.1093 | 1598.0821 |  | 1774.2038 | 1695.9964 | 1610.3128 | [5 empty] | Sc | 210.7414 | 4.0907 | 1.4087 | 1760.9023 | 1676.205 | 1598.1915 | 1774.3085 | 1696.1063 | 1610.4228
Ti | [14 empty] | Ti | 87.9524 | 4.5151 | 1.5074 | 1649.4202 | 1587.4759 | 1520.8253 |  |  | 1750.5068 | 1666.653 | 1587.2783 |  | 1763.8604 | 1685.1423 | 1599.4506 | [5 empty] | Ti | 191.1202 | 4.5107 | 1.5064 | 1750.611 | 1666.7487 | 1587.3876 | 1763.9651 | 1685.2521 | 1599.5606
V | [14 empty] | V | 80.1198 | 4.9565 | 1.6007 | 1640.685 | 1579.2008 | 1512.1498 |  |  | 1740.6848 | 1657.6279 | 1576.9671 |  | 1753.9886 | 1674.7831 | 1589.0837 | [5 empty] | V | 174.1 | 4.9516 | 1.5997 | 1740.789 | 1657.7237 | 1577.0764 | 1754.0933 | 1674.893 | 1589.1937
Cr | 235.3321 | 0.57273 | -0.55734 | 1771.6501 | 1710.9232 | 1639.8477 | 1575.971 | 1842.7886 | 1770.6623 | 1712.8727 | 1652.1529 | 1764.9751 | 1706.3136 | 1644.0747 | Cr | 73.2693 | 5.42 | 1.6901 | 1632.3144 | 1571.2709 | 1503.8363 |  |  | 1731.2726 | 1648.9794 | 1567.0861 |  | 1744.5287 | 1664.8561 | 1579.1494 | [5 empty] | Cr | 159.2139 | 5.4146 | 1.6891 | 1731.3768 | 1649.0751 | 1567.1954 | 1744.6334 | 1664.966 | 1579.2593
Mn | 211.688 | 0.6367 | -0.45145 | 1763.0938 | 1702.9255 | 1631.4436 | 1568.6668 | 1832.8725 | 1761.2683 | 1703.0242 | 1641.0977 | 1756.9582 | 1698.0419 | 1634.2635 | Mn | 67.2577 | 5.9044 | 1.7757 | 1624.2969 | 1563.6755 | 1495.8736 |  |  | 1722.2575 | 1640.6958 | 1557.622 |  | 1735.4679 | 1655.348 | 1569.6342 | [5 empty] | Mn | 146.1507 | 5.8986 | 1.7747 | 1722.3617 | 1640.7915 | 1557.7314 | 1735.5726 | 1655.4579 | 1569.7441
Fe | 191.2699 | 0.70467 | -0.35002 | 1754.8975 | 1695.2643 | 1623.3932 | 1561.67 | 1823.3736 | 1752.2696 | 1693.5902 | 1630.5078 | 1749.2786 | 1690.1183 | 1624.8653 | Fe | 61.9637 | 6.4089 | 1.8577 | 1616.619 | 1556.4019 | 1488.2482 |  |  | 1713.6243 | 1632.763 | 1548.5588 |  | 1726.791 | 1646.2426 | 1560.522 | [5 empty] | Fe | 134.6467 | 6.4025 | 1.8567 | 1713.7285 | 1632.8587 | 1548.6682 | 1726.8957 | 1646.3525 | 1560.632
Co | 173.7697 | 0.77564 | -0.25407 | 1747.1436 | 1688.0166 | 1615.7773 | 1555.0508 | 1814.3873 | 1743.7565 | 1684.6653 | 1620.4894 | 1742.0134 | 1682.6222 | 1615.9742 | Co | [20 empty] | Co | 124.4074 | 6.9295 | 1.9358 | 1705.3996 | 1625.2057 | 1539.9245 | 1718.5246 | 1637.5681 | 1551.8411
Ni | 158.2691 | 0.8516 | -0.16063 | 1739.5933 | 1680.9593 | 1608.3614 | 1548.6054 | 1805.6371 | 1735.467 | 1675.9749 | 1610.7341 | 1734.9391 | 1675.3231 | 1607.3166 | Ni | [20 empty] | Ni | 115.2906 | 7.4774 | 2.0119 | 1697.3854 | 1617.8417 | 1531.5111 | 1710.4698 | 1629.1156 | 1543.3823
Cu | 144.9949 | 0.92957 | -0.073037 | 1732.5146 | 1674.3428 | 1601.4087 | 1542.5627 | 1797.4334 | 1727.6953 | 1667.8273 | 1601.5881 | 1728.3067 | 1668.4799 | 1599.1998 | Cu | [20 empty] | Cu | 107.1251 | 8.0474 | 2.0853 | 1689.6499 | 1610.7339 | 1523.3904 | 1702.6951 | 1620.9571 | 1535.2177
Zn | 133.2463 | 1.0115 | 0.011463 | 1725.6864 | 1667.9603 | 1594.7019 | 1536.7336 | 1789.5199 | 1720.1985 | 1659.9679 | 1592.7657 | 1721.9089 | 1661.8788 | 1591.3702 | Zn | [20 empty] | Zn | 99.7966 | 8.6384 | 2.1562 | 1682.1877 | 1603.8771 | 1515.5566 | 1695.1951 | 1613.0868 | 1527.3416
Ga | 122.8099 | 1.0975 | 0.093024 | 1719.0955 | 1661.7997 | 1588.2283 | 1531.1073 | 1781.8815 | 1712.9624 | 1652.3817 | 1584.25 | 1715.7334 | 1655.5071 | 1583.8127 | Ga | [20 empty] | Ga | 93.1845 | 9.2513 | 2.2248 | 1674.9688 | 1597.2439 | 1507.9781 | 1687.9397 | 1605.4731 | 1519.7222
Ge | 113.5061 | 1.1874 | 0.17181 | 1712.7293 | 1655.8492 | 1581.9754 | 1525.6727 | 1774.5036 | 1705.9729 | 1645.0541 | 1576.0246 | 1709.7686 | 1649.3527 | 1576.513 | Ge | [20 empty] | Ge | 87.2085 | 9.8853 | 2.291 | 1667.9893 | 1590.8307 | 1500.651 | 1680.9247 | 1598.1118 | 1512.3555
As | 105.1835 | 1.2814 | 0.24796 | 1706.5757 | 1650.0974 | 1575.9314 | 1520.4197 | 1767.372 | 1699.2169 | 1637.9713 | 1568.0739 | 1704.0029 | 1643.4039 | 1569.4569 | As | [20 empty] | As | 81.7662 | 10.5432 | 2.3555 | 1661.2038 | 1584.5957 | 1493.5275 | 1674.1048 | 1590.9552 | 1505.1936
Se | 97.7848 | 1.3784 | 0.32089 | 1700.6818 | 1644.5882 | 1570.1423 | 1515.3882 | 1760.5414 | 1692.7459 | 1631.1873 | 1560.4587 | 1698.4805 | 1637.706 | 1562.6986 | Se | [20 empty] | Se | 76.8262 | 11.2212 | 2.4178 | 1654.6413 | 1578.5657 | 1486.6382 | 1667.5091 | 1584.0339 | 1498.267
Br | 91.1117 | 1.4793 | 0.39158 | 1694.9699 | 1639.2493 | 1564.5321 | 1510.5123 | 1753.9217 | 1686.4749 | 1624.6129 | 1553.0787 | 1693.1287 | 1632.1841 | 1556.1491 | Br | [20 empty] | Br | 72.3032 | 11.9231 | 2.4785 | 1648.2518 | 1572.6946 | 1479.9304 | 1661.0872 | 1577.2949 | 1491.523
Kr | 85.0222 | 1.5853 | 0.46075 | 1689.3802 | 1634.0245 | 1559.0418 | 1505.7405 | 1747.4436 | 1680.3378 | 1618.179 | 1545.8565 | 1687.8913 | 1626.7803 | 1549.7396 | Kr | 250.1597 | 1.5875 | 0.46213 | 1747.3143 | 1680.2153 | 1618.0506 |  |  | 1860.5816 | 1767.7967 | 1702.8353 |  | 1874.4929 | 1801.237 | 1715.6317 | [5 empty] | Kr
Rb | 79.6017 | 1.6932 | 0.52663 | 1684.0567 | 1629.0486 | 1553.8131 | 1501.1961 | 1741.2741 | 1674.4932 | 1612.0516 | 1538.9784 | 1682.9035 | 1621.6339 | 1543.6354 | Rb | 234.2109 | 1.6956 | 0.52801 | 1741.1448 | 1674.3707 | 1611.9232 |  |  | 1853.6444 | 1761.4224 | 1695.5526 |  | 1867.5206 | 1793.9205 | 1708.3097 | [5 empty] | Rb
Sr | 74.6238 | 1.8062 | 0.5912 | 1678.8385 | 1624.171 | 1548.6877 | 1496.7415 | 1735.2265 | 1668.764 | 1606.0454 | 1532.2362 | 1678.0142 | 1616.5893 | 1537.6519 | Sr | 219.5647 | 1.8087 | 0.59258 | 1735.0972 | 1668.6415 | 1605.917 |  |  | 1846.8444 | 1755.174 | 1688.4139 |  | 1860.6861 | 1786.7485 | 1701.1324 | [5 empty] | Sr
Y | 70.1223 | 1.9221 | 0.65342 | 1673.8107 | 1619.4715 | 1543.7494 | 1492.4495 | 1729.3997 | 1663.244 | 1600.2583 | 1525.74 | 1673.3033 | 1611.7287 | 1531.8868 | Y | 206.32 | 1.9248 | 0.6548 | 1729.2704 | 1663.1215 | 1600.1299 |  |  | 1840.2925 | 1749.1538 | 1681.5357 |  | 1854.101 | 1779.8383 | 1694.2171 | [5 empty] | Y
Zr | 66.0359 | 2.041 | 0.71346 | 1668.9587 | 1614.9363 | 1538.9837 | 1488.3075 | 1723.7765 | 1657.917 | 1594.6736 | 1519.471 | 1668.7572 | 1607.0381 | 1526.3232 | Zr | 194.2964 | 2.0439 | 0.71484 | 1723.6472 | 1657.7945 | 1594.5452 |  |  | 1833.9697 | 1743.344 | 1674.898 |  | 1847.7462 | 1773.1697 | 1687.5435 | [5 empty] | Zr
Nb | 62.2554 | 2.165 | 0.77242 | 1664.1948 | 1610.4834 | 1534.3047 | 1484.2408 | 1718.2555 | 1652.6867 | 1589.1904 | 1513.3159 | 1664.2936 | 1602.4327 | 1520.8607 | Nb | 183.1733 | 2.168 | 0.7738 | 1718.1262 | 1652.5642 | 1589.062 |  |  | 1827.7618 | 1737.6397 | 1668.3809 |  | 1841.5068 | 1766.6223 | 1680.9912 | [5 empty] | Nb
Mo | [14 empty] | Mo | 173.0281 | 2.2951 | 0.83078 | 1712.7901 | 1647.509 | 1583.7623 |  |  | 1821.7617 | 1732.1265 | 1662.0819 |  | 1835.4763 | 1760.294 | 1674.6583 | [5 empty] | Mo
Tc | [14 empty] | Tc | 163.6771 | 2.4262 | 0.88633 | 1707.587 | 1642.5799 | 1578.5947 |  |  | 1815.9112 | 1726.7507 | 1655.9401 |  | 1829.5962 | 1754.1236 | 1668.4832 | [5 empty] | Tc
Ru | [14 empty] | Ru | 155.0423 | 2.5613 | 0.94053 | 1702.5113 | 1637.7715 | 1573.5537 |  |  | 1810.204 | 1721.5065 | 1649.9486 |  | 1823.86 | 1748.1042 | 1662.4594 | [5 empty] | Ru
Rh | [14 empty] | Rh | 147.1091 | 2.6995 | 0.99305 | 1697.5924 | 1633.1116 | 1568.6684 |  |  | 1804.673 | 1716.4243 | 1644.1422 |  | 1818.3011 | 1742.2708 | 1656.6216 | [5 empty] | Rh
Pd | [14 empty] | Pd | 139.7511 | 2.8416 | 1.0444 | 1692.787 | 1628.5593 | 1563.8958 |  |  | 1799.2697 | 1711.4594 | 1638.4697 |  | 1812.8703 | 1736.572 | 1650.9185 | [5 empty] | Pd
Ag | [14 empty] | Ag | 132.9602 | 2.9867 | 1.0942 | 1688.122 | 1624.1399 | 1559.2627 |  |  | 1794.0242 | 1706.6395 | 1632.963 |  | 1807.5983 | 1731.0396 | 1645.3821 | [5 empty] | Ag
Cd | [14 empty] | Cd | 126.5965 | 3.1369 | 1.1432 | 1683.5288 | 1619.7885 | 1554.7009 |  |  | 1788.8595 | 1701.8939 | 1627.5411 |  | 1802.4074 | 1725.5924 | 1639.9309 | [5 empty] | Cd
In | [14 empty] | In | 120.7038 | 3.29 | 1.1909 | 1679.0649 | 1615.5597 | 1550.2675 |  |  | 1783.8402 | 1697.2818 | 1622.2718 |  | 1797.3626 | 1720.2986 | 1634.6331 | [5 empty] | In | 262.2887 | 3.2868 | 1.1899 | 1783.9443 | 1697.3775 | 1622.3811 | 1797.4673 | 1720.4085 | 1634.743
Sn | [14 empty] | Sn | 115.2013 | 3.4472 | 1.2375 | 1674.6953 | 1611.4202 | 1545.9277 |  |  | 1778.9269 | 1692.7671 | 1617.1137 |  | 1792.4244 | 1715.1166 | 1629.4472 | [5 empty] | Sn | 250.3319 | 3.4437 | 1.2366 | 1779.031 | 1692.8628 | 1617.2231 | 1792.5291 | 1715.2264 | 1629.5572
Sb | [14 empty] | Sb | 110.0564 | 3.6083 | 1.2832 | 1670.4165 | 1607.3667 | 1541.6782 |  |  | 1774.1157 | 1688.3463 | 1612.063 |  | 1787.5889 | 1710.0423 | 1624.3692 | [5 empty] | Sb | 239.152 | 3.6047 | 1.2822 | 1774.2199 | 1688.442 | 1612.1723 | 1787.6936 | 1710.1522 | 1624.4791
Te | [14 empty] | Te | 105.2675 | 3.7724 | 1.3277 | 1666.2502 | 1603.4197 | 1537.5403 |  |  | 1769.4309 | 1684.0417 | 1607.1449 |  | 1782.8804 | 1705.1013 | 1619.4245 | [5 empty] | Te | 228.7458 | 3.7687 | 1.3267 | 1769.5351 | 1684.1374 | 1607.2542 | 1782.9851 | 1705.2112 | 1619.5345
I | [14 empty] | I | 100.75 | 3.9416 | 1.3716 | 1662.1423 | 1599.5282 | 1533.4605 |  |  | 1764.812 | 1679.7975 | 1602.2959 |  | 1778.238 | 1700.2297 | 1614.5493 | [5 empty] | I | 218.9292 | 3.9377 | 1.3706 | 1764.9161 | 1679.8932 | 1602.4052 | 1778.3427 | 1700.3396 | 1614.6592
Xe | [14 empty] | Xe | 96.5337 | 4.1138 | 1.4143 | 1658.1387 | 1595.7354 | 1529.4842 |  |  | 1760.3102 | 1675.661 | 1597.5699 |  | 1773.7134 | 1695.4818 | 1609.7978 | [5 empty] | Xe | 209.7672 | 4.1097 | 1.4133 | 1760.4144 | 1675.7567 | 1597.6792 | 1773.8181 | 1695.5916 | 1609.9078
Cs | [14 empty] | Cs | 92.548 | 4.2909 | 1.4565 | 1654.19 | 1591.9946 | 1525.5625 |  |  | 1755.8701 | 1671.5812 | 1592.9087 |  | 1769.2509 | 1690.7989 | 1605.1114 | [5 empty] | Cs | 201.1064 | 4.2867 | 1.4555 | 1755.9743 | 1671.6769 | 1593.0181 | 1769.3556 | 1690.9088 | 1605.2214
Ba | [14 empty] | Ba | 88.8386 | 4.4701 | 1.4974 | 1650.3591 | 1588.3654 | 1521.7578 |  |  | 1751.5626 | 1667.6231 | 1588.3866 |  | 1764.9215 | 1686.2557 | 1600.5649 | [5 empty] | Ba | 193.0459 | 4.4657 | 1.4964 | 1751.6667 | 1667.7188 | 1588.496 | 1765.0262 | 1686.3656 | 1600.6749
La | 161.8791 | 0.83261 | -0.18319 | 1741.4158 | 1682.6628 | 1610.1514 | 1550.1612 | 1807.7492 | 1737.4679 | 1678.0726 | 1613.0888 | 1736.6467 | 1677.085 | 1609.4064 | La | 85.305 | 4.6553 | 1.538 | 1646.5578 | 1584.7643 | 1517.9825 |  |  | 1747.2884 | 1663.6957 | 1583.8995 |  | 1760.6256 | 1681.7478 | 1596.0536 | [5 empty] | La | 185.3673 | 4.6507 | 1.537 | 1747.3925 | 1663.7914 | 1584.0089 | 1760.7303 | 1681.8576 | 1596.1636
Ce | 152.7127 | 0.88259 | -0.1249 | 1736.7053 | 1678.2598 | 1605.5248 | 1546.1401 | 1802.2901 | 1732.2963 | 1672.6508 | 1607.0027 | 1732.2332 | 1672.5312 | 1604.0051 | Ce | 81.9738 | 4.8444 | 1.5778 | 1642.8275 | 1581.2304 | 1514.2776 |  |  | 1743.0938 | 1659.8415 | 1579.4961 |  | 1756.4099 | 1677.3239 | 1591.6264 | [5 empty] | Ce | 178.1287 | 4.8396 | 1.5768 | 1743.198 | 1659.9372 | 1579.6054 | 1756.5145 | 1677.4337 | 1591.7363
Pr | 145.151 | 0.92857 | -0.074113 | 1732.6016 | 1674.424 | 1601.4941 | 1542.6369 | 1797.5341 | 1727.7908 | 1667.9273 | 1601.7005 | 1728.3881 | 1668.564 | 1599.2995 | Pr | 78.8147 | 5.0386 | 1.6171 | 1639.147 | 1577.7437 | 1510.6223 |  |  | 1738.9554 | 1656.0388 | 1575.1515 |  | 1752.2504 | 1672.9591 | 1587.2584 | [5 empty] | Pr | 171.264 | 5.0336 | 1.6161 | 1739.0595 | 1656.1345 | 1575.2609 | 1752.3551 | 1673.0689 | 1587.3683
Nd | 137.8786 | 0.97754 | -0.022712 | 1728.4479 | 1670.5416 | 1597.4143 | 1539.0911 | 1792.7204 | 1723.2305 | 1663.1465 | 1596.3338 | 1724.4963 | 1664.5485 | 1594.5368 | Nd | 75.8611 | 5.2348 | 1.6553 | 1635.5698 | 1574.3549 | 1507.0695 |  |  | 1734.9331 | 1652.3429 | 1570.9289 |  | 1748.2078 | 1668.7168 | 1583.013 | [5 empty] | Nd | 164.8457 | 5.2296 | 1.6543 | 1735.0373 | 1652.4386 | 1571.0383 | 1748.3125 | 1668.8267 | 1583.1229
Pm | 131.045 | 1.0285 | 0.028121 | 1724.3402 | 1666.702 | 1593.3797 | 1535.5845 | 1787.9598 | 1718.7206 | 1658.4184 | 1591.0264 | 1720.6475 | 1660.5774 | 1589.8266 | Pm | 73.0266 | 5.438 | 1.6934 | 1632.0036 | 1570.9764 | 1503.5276 |  |  | 1730.9231 | 1648.6583 | 1566.7192 |  | 1744.1775 | 1664.4875 | 1578.7805 | [5 empty] | Pm | 158.6864 | 5.4326 | 1.6924 | 1731.0273 | 1648.754 | 1566.8286 | 1744.2822 | 1664.5974 | 1578.8905
Sm | 124.7412 | 1.0805 | 0.077421 | 1720.3564 | 1662.9783 | 1589.4668 | 1532.1837 | 1783.3428 | 1714.3468 | 1653.833 | 1585.8792 | 1716.9149 | 1656.7261 | 1585.2586 | Sm | 70.3963 | 5.6412 | 1.7301 | 1628.5681 | 1567.7219 | 1500.1157 |  |  | 1727.0603 | 1645.1088 | 1562.664 |  | 1740.295 | 1660.4134 | 1574.7034 | [5 empty] | Sm | 152.9707 | 5.6356 | 1.7291 | 1727.1644 | 1645.2045 | 1562.7733 | 1740.3997 | 1660.5233 | 1574.8133
Eu | 119.2266 | 1.1305 | 0.12264 | 1716.7026 | 1659.5631 | 1585.878 | 1529.0646 | 1779.1083 | 1710.3352 | 1649.6274 | 1581.1583 | 1713.4914 | 1653.1939 | 1581.0689 | Eu | 67.8675 | 5.8513 | 1.7667 | 1625.1421 | 1564.4762 | 1496.713 |  |  | 1723.2079 | 1641.569 | 1558.6197 |  | 1736.4231 | 1656.3504 | 1570.6373 | [5 empty] | Eu | 147.4757 | 5.8456 | 1.7657 | 1723.3121 | 1641.6648 | 1558.7291 | 1736.5278 | 1656.4602 | 1570.7473
Gd | 113.7935 | 1.1844 | 0.16928 | 1712.9336 | 1656.0402 | 1582.1761 | 1525.8472 | 1774.7404 | 1706.1973 | 1645.2893 | 1576.2886 | 1709.96 | 1649.5503 | 1576.7472 | Gd | 65.5033 | 6.0625 | 1.8021 | 1621.8215 | 1561.3305 | 1493.4151 |  |  | 1719.4741 | 1638.1382 | 1554.7 |  | 1732.6705 | 1652.4124 | 1566.6964 | [5 empty] | Gd | 142.3383 | 6.0565 | 1.8011 | 1719.5783 | 1638.2339 | 1554.8094 | 1732.7752 | 1652.5223 | 1566.8064
Tb | 108.7462 | 1.2394 | 0.21465 | 1709.2675 | 1652.6134 | 1578.5752 | 1522.7175 | 1770.4916 | 1702.1722 | 1641.0695 | 1571.5518 | 1706.525 | 1646.0061 | 1572.5434 | Tb | 63.2478 | 6.2787 | 1.8372 | 1618.5399 | 1558.2217 | 1490.156 |  |  | 1715.7843 | 1634.7477 | 1550.8264 |  | 1728.9619 | 1648.5207 | 1562.8019 | [5 empty] | Tb | 137.4371 | 6.2725 | 1.8362 | 1715.8884 | 1634.8435 | 1550.9357 | 1729.0666 | 1648.6306 | 1562.9118
Dy | 104.2887 | 1.2924 | 0.2565 | 1705.8853 | 1649.452 | 1575.2532 | 1519.8303 | 1766.5719 | 1698.4589 | 1637.1766 | 1567.1819 | 1703.356 | 1642.7364 | 1568.6653 | Dy | 61.086 | 6.5009 | 1.8719 | 1615.2829 | 1555.1362 | 1486.9213 |  |  | 1712.122 | 1631.3826 | 1546.9817 |  | 1725.2811 | 1644.6582 | 1558.9364 | [5 empty] | Dy | 132.7395 | 6.4945 | 1.871 | 1712.2262 | 1631.4783 | 1547.0911 | 1725.3858 | 1644.768 | 1559.0464
Ho | 100.0336 | 1.3474 | 0.29816 | 1702.5191 | 1646.3056 | 1571.9469 | 1516.9567 | 1762.6707 | 1694.7631 | 1633.3021 | 1562.8326 | 1700.202 | 1639.4822 | 1564.8054 | Ho | [20 empty] | Ho | 128.2951 | 6.7195 | 1.905 | 1708.64 | 1628.1831 | 1543.3263 | 1721.7814 | 1640.9857 | 1555.2613
Er | 95.907 | 1.4053 | 0.34028 | 1699.1149 | 1643.1236 | 1568.6033 | 1514.0506 | 1758.7254 | 1691.0256 | 1629.3838 | 1558.4342 | 1697.0124 | 1636.1912 | 1560.9019 | Er | [20 empty] | Er | 124.0672 | 6.9485 | 1.9385 | 1705.1113 | 1624.9407 | 1539.6218 | 1718.2348 | 1637.264 | 1551.5368
Tm | 92.2334 | 1.4613 | 0.37934 | 1695.9588 | 1640.1736 | 1565.5034 | 1511.3564 | 1755.0677 | 1687.5605 | 1625.7511 | 1554.3564 | 1694.0552 | 1633.1401 | 1557.283 | Tm | [20 empty] | Tm | 120.0756 | 7.1795 | 1.9712 | 1701.6677 | 1621.7765 | 1536.0067 | 1714.7738 | 1633.6321 | 1547.9022
Yb | 88.6557 | 1.5203 | 0.4189 | 1692.7618 | 1637.1853 | 1562.3632 | 1508.6273 | 1751.3626 | 1684.0505 | 1622.0713 | 1550.2257 | 1691.0598 | 1630.0494 | 1553.6171 | Yb | 260.8503 | 1.5224 | 0.42028 | 1751.2333 | 1683.928 | 1621.9429 |  |  | 1864.9883 | 1771.8458 | 1707.4614 |  | 1878.9219 | 1805.8847 | 1720.2828 | [5 empty] | Yb | 116.2545 | 7.4154 | 2.0036 | 1698.2621 | 1618.6473 | 1532.4315 | 1711.3509 | 1630.0403 | 1544.3077
Lu | 85.2911 | 1.5803 | 0.45759 | 1689.6353 | 1634.263 | 1559.2924 | 1505.9583 | 1747.7393 | 1680.618 | 1618.4727 | 1546.1862 | 1688.1304 | 1627.027 | 1550.0322 | Lu | 250.9509 | 1.5824 | 0.45897 | 1747.61 | 1680.4955 | 1618.3443 |  |  | 1860.9141 | 1768.1022 | 1703.1843 |  | 1874.8271 | 1801.5877 | 1715.9826 | [5 empty] | Lu | 112.6101 | 7.6554 | 2.0354 | 1694.9082 | 1615.5655 | 1528.9106 | 1707.98 | 1626.5029 | 1540.7677
Hf | 81.9728 | 1.6442 | 0.49727 | 1686.4286 | 1631.2656 | 1556.1428 | 1503.2209 | 1744.023 | 1677.0973 | 1614.7817 | 1542.043 | 1685.1258 | 1623.9269 | 1546.3552 | Hf | 241.1874 | 1.6465 | 0.49866 | 1743.8937 | 1676.9748 | 1614.6533 |  |  | 1856.7354 | 1764.2625 | 1698.7975 |  | 1870.6272 | 1797.1804 | 1711.5721 | [5 empty] | Hf | 109.1459 | 7.8984 | 2.0667 | 1691.6178 | 1612.5421 | 1525.4563 | 1704.673 | 1623.0326 | 1537.2948
Ta | 78.8569 | 1.7092 | 0.53603 | 1683.2971 | 1628.3385 | 1553.067 | 1500.5476 | 1740.3937 | 1673.6591 | 1611.1773 | 1537.9968 | 1682.1917 | 1620.8995 | 1542.7644 | Ta | 232.0195 | 1.7116 | 0.53741 | 1740.2644 | 1673.5366 | 1611.0488 |  |  | 1852.6545 | 1760.5127 | 1694.5134 |  | 1866.5257 | 1792.8764 | 1707.2648 | [5 empty] | Ta | 105.8364 | 8.1454 | 2.0975 | 1688.3754 | 1609.5627 | 1522.0524 | 1701.4141 | 1619.6129 | 1533.8725
W | 75.9692 | 1.7742 | 0.57333 | 1680.2823 | 1625.5206 | 1550.1059 | 1497.974 | 1736.8998 | 1670.3492 | 1607.7073 | 1534.1017 | 1679.367 | 1617.9851 | 1539.3075 | W | 223.523 | 1.7766 | 0.57472 | 1736.7705 | 1670.2268 | 1607.5789 |  |  | 1848.7259 | 1756.9029 | 1690.3891 |  | 1862.5772 | 1788.7329 | 1703.1183 | [5 empty] | W | 102.6605 | 8.3974 | 2.1279 | 1685.1671 | 1606.6148 | 1518.6844 | 1698.1896 | 1616.2291 | 1530.4863
Re | 73.1662 | 1.8421 | 0.61093 | 1677.2444 | 1622.681 | 1547.122 | 1495.3807 | 1733.3791 | 1667.0139 | 1604.2106 | 1530.1765 | 1676.5206 | 1615.0482 | 1535.8241 | Re | 215.2758 | 1.8447 | 0.61231 | 1733.2498 | 1666.8914 | 1604.0822 |  |  | 1844.7671 | 1753.2653 | 1686.2331 |  | 1858.5983 | 1784.5576 | 1698.9399 | [5 empty] | Re | 99.6467 | 8.6514 | 2.1577 | 1682.0294 | 1603.7316 | 1515.3904 | 1695.036 | 1612.9198 | 1527.1745
Os | 70.5996 | 1.9091 | 0.64664 | 1674.3588 | 1619.9839 | 1544.2878 | 1492.9174 | 1730.0349 | 1663.8458 | 1600.8893 | 1526.4483 | 1673.8169 | 1612.2586 | 1532.5153 | Os | 207.7243 | 1.9117 | 0.64802 | 1729.9056 | 1663.7233 | 1600.7608 |  |  | 1841.0068 | 1749.8101 | 1682.2856 |  | 1854.819 | 1780.5917 | 1694.971 | [5 empty] | Os | 96.7396 | 8.9113 | 2.1873 | 1678.9115 | 1600.8667 | 1512.1172 | 1691.9023 | 1609.6314 | 1523.8836
Ir | [14 empty] | Ir | 200.3805 | 1.9818 | 0.68401 | 1726.5348 | 1660.53 | 1597.413 |  |  | 1837.2166 | 1746.3274 | 1678.3065 |  | 1851.0095 | 1776.5941 | 1690.9705 | [5 empty] | Ir | 93.9665 | 9.1743 | 2.2164 | 1675.8489 | 1598.0526 | 1508.902 | 1688.8242 | 1606.4013 | 1520.6511
Pt | [14 empty] | Pt | 193.4438 | 2.0529 | 0.71924 | 1723.2354 | 1657.4043 | 1594.1362 |  |  | 1833.5066 | 1742.9185 | 1674.4118 |  | 1847.2808 | 1772.6813 | 1687.0547 | [5 empty] | Pt | 91.3094 | 9.4413 | 2.2451 | 1672.8282 | 1595.277 | 1505.7309 | 1685.7882 | 1603.2154 | 1517.4628
Au | [14 empty] | Au | 186.8833 | 2.1249 | 0.75374 | 1720.0041 | 1654.3432 | 1590.927 |  |  | 1829.8733 | 1739.58 | 1670.5976 |  | 1843.6291 | 1768.8493 | 1683.2199 | [5 empty] | Au | 88.7618 | 9.7123 | 2.2734 | 1669.8484 | 1592.5389 | 1502.6027 | 1682.7932 | 1600.0726 | 1514.3177
Hg | [14 empty] | Hg | 180.7532 | 2.197 | 0.7871 | 1716.8807 | 1651.3842 | 1587.8249 |  |  | 1826.3613 | 1736.3528 | 1666.9106 |  | 1840.0992 | 1765.1452 | 1679.513 | [5 empty] | Hg | 86.3092 | 9.9883 | 2.3014 | 1666.8978 | 1589.8278 | 1499.5052 | 1679.8277 | 1596.9607 | 1511.2035
Tl | [14 empty] | Tl | 174.7042 | 2.2731 | 0.82113 | 1713.693 | 1648.3644 | 1584.659 |  |  | 1822.7769 | 1733.0593 | 1663.1477 |  | 1836.4967 | 1761.3648 | 1675.7298 | [5 empty] | Tl | 83.9559 | 10.2682 | 2.3291 | 1663.9867 | 1587.1529 | 1496.449 | 1676.9018 | 1593.8903 | 1508.1308
Pb | [14 empty] | Pb | 169.1191 | 2.3481 | 0.85363 | 1710.6501 | 1645.4817 | 1581.6369 |  |  | 1819.3554 | 1729.9154 | 1659.5558 |  | 1833.0578 | 1757.7562 | 1672.1185 | [5 empty] | Pb | 81.7042 | 10.5512 | 2.3562 | 1661.1239 | 1584.5224 | 1493.4437 | 1674.0246 | 1590.871 | 1505.1093
Bi | [14 empty] | Bi | 163.7447 | 2.4252 | 0.88592 | 1707.6256 | 1642.6165 | 1578.6331 |  |  | 1815.9546 | 1726.7906 | 1655.9857 |  | 1829.6398 | 1754.1694 | 1668.5291 | [5 empty] | Bi | 79.5408 | 10.8382 | 2.3831 | 1658.298 | 1581.9258 | 1490.4771 | 1671.1844 | 1587.8906 | 1502.1266
Po | [14 empty] | Po | 158.6379 | 2.5033 | 0.91761 | 1704.6584 | 1639.8055 | 1575.6861 |  |  | 1812.6182 | 1723.7248 | 1652.483 |  | 1826.2864 | 1750.6505 | 1665.0075 | [5 empty] | Po | 77.4542 | 11.1302 | 2.4097 | 1655.4987 | 1579.3536 | 1487.5383 | 1668.3709 | 1584.9381 | 1499.172
At | [14 empty] | At | 153.7804 | 2.5824 | 0.9487 | 1701.7459 | 1637.0464 | 1572.7936 |  |  | 1809.3433 | 1720.7157 | 1649.0451 |  | 1822.995 | 1747.1965 | 1661.551 | [5 empty] | At | 75.4479 | 11.4261 | 2.4359 | 1652.735 | 1576.8141 | 1484.637 | 1665.5932 | 1582.0233 | 1496.255
Rn | [14 empty] | Rn | 149.2115 | 2.6614 | 0.97886 | 1698.9213 | 1634.3706 | 1569.9883 |  |  | 1806.1673 | 1717.7974 | 1645.7109 |  | 1819.8029 | 1743.8468 | 1658.1988 | [5 empty] | Rn | 73.5178 | 11.7261 | 2.4618 | 1650.0061 | 1574.3066 | 1481.7721 | 1662.8504 | 1579.1451 | 1493.3747
Fr | [14 empty] | Fr | 144.7477 | 2.7435 | 1.0092 | 1696.0769 | 1631.6759 | 1567.1632 |  |  | 1802.9689 | 1714.8585 | 1642.3532 |  | 1816.5883 | 1740.4735 | 1654.823 | [5 empty] | Fr | 71.6601 | 12.0301 | 2.4874 | 1647.311 | 1571.8302 | 1478.9429 | 1660.1417 | 1576.3027 | 1490.5301
Ra | [14 empty] | Ra | 140.4934 | 2.8266 | 1.0391 | 1693.2831 | 1629.0293 | 1564.3886 |  |  | 1799.8275 | 1711.972 | 1639.0554 |  | 1813.431 | 1737.1603 | 1651.5073 | [5 empty] | Ra | 69.8657 | 12.3391 | 2.5128 | 1644.6406 | 1569.3764 | 1476.1394 | 1657.4577 | 1573.4862 | 1487.7115
Ac | [14 empty] | Ac | 136.3882 | 2.9117 | 1.0687 | 1690.5059 | 1626.3983 | 1561.6303 |  |  | 1796.7047 | 1709.1026 | 1635.777 |  | 1810.2924 | 1733.8667 | 1648.2113 | [5 empty] | Ac | 68.1428 | 12.6511 | 2.5377 | 1642.0112 | 1566.9604 | 1473.3791 | 1654.815 | 1570.713 | 1484.9363
Th | [14 empty] | Th | 132.4277 | 2.9987 | 1.0982 | 1687.7461 | 1623.7838 | 1558.8894 |  |  | 1793.6016 | 1706.2512 | 1632.5193 |  | 1807.1735 | 1730.5938 | 1644.936 | [5 empty] | Th | 66.4772 | 12.968 | 2.5625 | 1639.4053 | 1564.5659 | 1470.6434 | 1652.1959 | 1567.9646 | 1482.1858
Pa | [14 empty] | Pa | 128.7324 | 3.0848 | 1.1265 | 1685.0957 | 1621.273 | 1556.2571 |  |  | 1790.6214 | 1703.5128 | 1629.3907 |  | 1804.1782 | 1727.4507 | 1641.7905 | [5 empty] | Pa | 64.8667 | 13.29 | 2.587 | 1636.8227 | 1562.1929 | 1467.9322 | 1649.6002 | 1565.2407 | 1479.46
U | [14 empty] | U | 125.1193 | 3.1739 | 1.155 | 1682.4296 | 1618.7472 | 1553.6092 |  |  | 1787.6236 | 1700.7582 | 1626.2436 |  | 1801.1652 | 1724.2889 | 1638.6264 | [5 empty] | U | 63.323 | 13.614 | 2.6111 | 1634.2864 | 1559.8624 | 1465.2696 | 1647.0511 | 1562.5658 | 1476.783
Np | [14 empty] | Np | 121.9654 | 3.256 | 1.1805 | 1680.0386 | 1616.4822 | 1551.2346 |  |  | 1784.9351 | 1698.2879 | 1623.4212 |  | 1798.4631 | 1721.4534 | 1635.7887 | [5 empty] | Np | 61.829 | 13.943 | 2.635 | 1631.7721 | 1557.552 | 1462.63 | 1644.524 | 1559.9139 | 1474.1292
Pu | [14 empty] | Pu | 118.6109 | 3.3481 | 1.2084 | 1677.4268 | 1614.0078 | 1548.6406 |  |  | 1781.9982 | 1695.5893 | 1620.3381 |  | 1795.5114 | 1718.3559 | 1632.689 | [5 empty] | Pu | 60.3784 | 14.2779 | 2.6587 | 1629.2721 | 1555.2549 | 1460.0055 | 1642.0114 | 1557.2772 | 1471.4905
Am | [14 empty] | Am | [20 empty] | Am
Cm | [14 empty] | Cm | [20 empty] | Cm
Bk | [14 empty] | Bk | [20 empty] | Bk
Cf | [14 empty] | Cf | [20 empty] | Cf
Es | [14 empty] | Es | [20 empty] | Es

## Positions
Calculated peak positions for chosen crystal (enter in 2d) | [12 empty] | LDE1 | 60
Jeol microprobe with R=140mm | [12 empty] | LDE2 | 98
Position calculated from L,mm = (R/d)*n*lamda | [12 empty] | LDEB | 145
Lamda = 12.39854141/E | [12 empty] | LDEC | 200
Radius | 140 | Low limit | 70 | High limit | 260
2d | 60 | <--Enter 2d value here
Order | 1
Lamda conv | 12398.5414 | 12398.5414
Symbol | Ka | Kb | La1 | Lb1 | Lb2 | Lb3 | Lb4 | Lg1 | Lg2,3 | Ma | Mb | Mg | Crystal | 2d
H | [12 empty] | LiF420 | 1.802
He | [12 empty] | LiF220 | 2.848
Li | [12 empty] | LIF | 4.0267
Be | 578.5986 | 578.5986 | [10 empty] | LIFH | 4.0267
B | 316.1741 | 316.1741 | [10 empty] | NaCl | 5.641
C | 208.8804 | 208.8804 | [10 empty] | Si | 6.2709
N | 147.6017 | 147.6017 | [10 empty] | Ge | 6.532
O | 110.2093 | 110.2093 | [10 empty] | QTZ011 | 6.686
F | 85.4651 | 85.4651 | [10 empty] | QTZ010 | 8.5
Ne | 68.1506 | 67.4357 | [10 empty] | PET | 8.75
Na | 55.581 | 54.0241 | [10 empty] | PETH | 8.75
Mg | 46.1402 | 44.4392 | [10 empty] | PETJ | 8.75
Al | 38.9105 | 37.1611 | [10 empty] | EDT | 8.808
Si | 33.2528 | 31.5141 | [10 empty] | ADP | 10.648
P | 28.7288 | 27.05 | [10 empty] | GYPSUM | 15.19
S | 25.0693 | 23.4821 | [10 empty] | MICA | 19.8
Cl | 22.0671 | 20.5468 | [10 empty] | TAP | 25.745
Ar | 19.5605 | 18.1322 | [10 empty] | RAP | 26.12
K | 17.4592 | 16.117 | [10 empty] | KAP | 26.63
Ca | 15.6717 | 14.4181 | 169.677 | 169.677 | 169.677 | 169.677 | 169.677 | 169.677 | 169.677 |  |  |  | WSi55 | 55
Sc | 14.1432 | 12.9702 | 146.4807 | 144.6496 | 146.4807 | 144.6496 | 144.6496 | 144.6496 | 144.6496 |  |  |  | LDE1 | 60
Ti | 12.8264 | 11.7315 | 128.0085 | 126.3316 | 128.0085 | 126.3316 | 126.3316 | 126.3316 | 126.3316 |  |  |  | WSi60 | 61
V | 11.6841 | 10.6615 | 113.2287 | 111.4834 | 113.2287 | 111.4834 | 111.4834 | 111.4834 | 111.4834 |  |  |  | LDE1 | 61.1
Cr | 10.6851 | 9.7293 | 100.9771 | 99.245 | 100.9771 | 98.9057 | 98.9057 | 99.245 | 98.9057 |  |  |  | WSi63 | 63
Mn | 9.8084 | 8.9152 | 90.8318 | 89.1523 | 90.8318 | 88.4707 | 88.4707 | 89.1523 | 88.4707 |  |  |  | NiCrBN | 65
Fe | 9.0364 | 8.1978 | 82.0707 | 80.4727 | 82.0707 | 80.2495 | 80.2495 | 80.4727 | 80.2495 |  |  |  | LSM | 80
Co | 8.3492 | 7.5644 | 74.5617 | 73.1477 | 74.5617 | 73.0554 | 73.0554 | 73.1477 | 73.0554 |  |  |  | NiC | 80.2
Ni | 7.7373 | 7.0006 | 67.9106 | 66.5821 | 67.9106 | 61.4876 | 61.4876 | 66.5821 | 61.4876 |  |  |  | OV95 | 96.8
Cu | 7.1893 | 6.4975 | 62.2149 | 60.9051 | 62.2149 | 56.559 | 56.559 | 60.9051 | 56.559 |  |  |  | STE | 98
Zn | 6.6975 | 6.0447 | 57.1738 | 55.9032 | 57.1738 | 52.2673 | 52.2673 | 55.9032 | 52.2673 |  |  |  | LDE2 | 98
Ga | 6.2538 | 5.6372 | 52.6957 | 51.431 | 52.6957 | 48.3374 | 48.3374 | 51.431 | 48.3374 |  |  |  | PbSd | 100.6
Ge | 5.8527 | 5.2686 | 48.7036 | 47.465 | 48.7036 | 44.714 | 44.9921 | 47.465 | 44.714 |  |  |  | PbSTEA | 100.6
As | 5.4875 | 4.9343 | 45.1325 | 43.9331 | 45.1325 | 41.6858 | 41.6858 | 43.9331 | 41.6858 |  |  |  | WC144 | 144
Se | 5.1559 | 4.6303 | 41.9578 | 40.7751 | 41.9578 | 38.8321 | 38.8321 | 40.7751 | 38.8321 |  |  |  | LDEB | 145
Br | 4.8524 | 4.3533 | 39.0945 | 37.916 | 39.0945 | 36.253 | 36.253 | 37.916 | 36.253 |  |  |  | OV160 | 160
Kr | 4.5743 | 4.1 | 36.4816 | 35.3451 | 36.4816 | 34.0954 | 34.0954 | 35.3451 | 33.8956 |  |  |  | MoB4C | 194.57
Rb | 4.3195 | 3.8674 | 34.1558 | 33.025 | 34.1558 | 31.6693 | 31.8261 | 33.025 | 28.2106 |  |  |  | OV200 | 197.4
Sr | 4.0847 | 3.6537 | 32.0198 | 30.908 | 32.0198 | 29.7174 | 29.8863 | 30.908 | 26.3358 |  |  |  | LDEC | 200
Y | 3.8682 | 3.4568 | 30.0883 | 28.9879 | 30.0883 | 27.9246 | 28.0873 | 28.9879 | 24.6527
Zr | 3.6678 | 3.2748 | 28.3349 | 27.241 | 28.3349 | 26.288 | 26.4563 | 27.241 | 23.1162
Nb | 3.4824 | 3.1069 | 26.7128 | 25.6357 | 24.4444 | 24.7794 | 24.9503 | 23.5012 | 21.7192
Mo | 3.3103 | 2.9508 | 25.2333 | 24.1586 | 22.9785 | 23.3966 | 23.5586 | 22.0503 | 20.438
Tc | 3.1502 | 2.8061 | 23.8696 | 22.8064 | 21.9999 | 22.2111 | 22.3225 | 20.7234 | 19.2866
Ru | 3.0012 | 2.6716 | 22.6103 | 21.5654 | 20.4019 | 20.941 | 21.109 | 19.5143 | 18.1892
Rh | 2.8621 | 2.5462 | 21.4534 | 20.4163 | 19.2802 | 19.8422 | 20.0138 | 18.4033 | 17.1997
Pd | 2.7322 | 2.4291 | 20.3804 | 19.3511 | 18.2408 | 18.8285 | 19.0016 | 17.3806 | 16.2848
Ag | 2.6107 | 2.3198 | 19.39 | 18.3624 | 17.2819 | 17.8911 | 18.0643 | 16.4375 | 15.4458
Cd | 2.4968 | 2.2172 | 18.462 | 17.4434 | 16.4002 | 17.0126 | 17.1844 | 15.5663 | 14.6444
In | 2.3899 | 2.1213 | 17.6026 | 16.593 | 15.5789 | 16.1936 | 16.3677 | 14.7564 | 13.9053
Sn | 2.2896 | 2.0312 | 16.8002 | 15.7958 | 14.8169 | 15.4293 | 15.6041 | 14.0063 | 13.2191
Sb | 2.1951 | 1.9464 | 16.0499 | 15.052 | 14.1087 | 14.7114 | 14.8893 | 13.3072 | 12.5782
Te | 2.1061 | 1.8667 | 15.3515 | 14.3573 | 13.4495 | 14.0437 | 14.2162 | 12.658 | 11.9817
I | 2.0222 | 1.7916 | 14.6927 | 13.7076 | 12.8349 | 13.4152 | 13.5885 | 12.0516 | 11.4212
Xe | 1.943 | 1.7208 | 14.0778 | 13.0994 | 12.261 | 12.8207 | 13.5885 | 11.487 | 10.9067
Cs | 1.8681 | 1.6538 | 13.4966 | 12.5238 | 11.722 | 12.2662 | 12.4457 | 10.9583 | 10.4327
Ba | 1.7972 | 1.5905 | 12.9556 | 11.9842 | 11.2197 | 11.7434 | 11.925 | 10.461 | 9.969
La | 1.7302 | 1.5306 | 12.4403 | 11.4756 | 10.7466 | 11.2502 | 11.4302 | 9.9948 | 9.5368 | 69.4596 | 67.7516 | 56.3387
Ce | 1.6665 | 1.4739 | 11.9545 | 10.9958 | 10.3082 | 10.7847 | 10.9645 | 9.5605 | 9.1377 | 65.5265 | 64.1462 | 53.8231
Pr | 1.6061 | 1.4199 | 11.4938 | 10.5411 | 9.8906 | 10.3469 | 10.5238 | 9.1521 | 8.7574 | 62.2819 | 60.9051 | 51.3397
Nd | 1.5487 | 1.3688 | 11.0631 | 10.1118 | 9.5024 | 9.9262 | 10.1118 | 8.764 | 8.3952 | 59.1614 | 58.034 | 49.0338
Pm | 1.4941 | 1.3202 | 10.6497 | 9.7064 | 9.1276 | 9.5305 | 9.7129 | 8.3952 | 8.0529 | 56.2292 | 55.1572 | 46.7743
Sm | 1.4422 | 1.2741 | 10.2661 | 9.3247 | 8.7839 | 9.1579 | 9.3383 | 8.0607 | 7.7188 | 53.5244 | 52.5999 | 44.8179
Eu | 1.3928 | 1.2301 | 9.8973 | 8.9622 | 8.4553 | 8.8053 | 8.9858 | 7.7353 | 7.4351 | 51.1581 | 50.182 | 42.9865
Gd | 1.3456 | 1.1882 | 9.5526 | 8.6191 | 8.1458 | 8.4702 | 8.6526 | 7.4313 | 7.1467 | 48.8269 | 47.8576 | 41.2695
Tb | 1.3007 | 1.1519 | 9.2236 | 8.2918 | 7.8539 | 8.1539 | 8.3372 | 7.1414 | 6.8799 | 46.6612 | 45.7029 | 39.6029
Dy | 1.2579 | 1.1136 | 8.9084 | 7.9829 | 7.5772 | 7.8507 | 8.0316 | 6.8725 | 6.6254 | 44.7485 | 43.6678 | 38.0157
Ho | 1.2169 | 1.0739 | 8.6101 | 7.689 | 7.3138 | 7.5614 | 7.7446 | 6.6148 | 6.38 | 42.9227 | 41.8365 | 36.7131
Er | 1.1777 | 1.0391 | 8.3264 | 7.4075 | 7.0656 | 7.2881 | 7.4706 | 6.3659 | 6.1442 | 41.1521 | 40.0969 | 35.216
Tm | 1.1403 | 1.006 | 8.0585 | 7.1423 | 6.8328 | 7.0295 | 7.2091 | 6.1383 | 5.9319 | 39.5758 | 38.4962 | 33.9354
Yb | 1.1044 | 0.97456 | 7.802 | 6.8864 | 6.6058 | 6.7775 | 6.9602 | 5.8622 | 5.7767 | 38.0407 | 36.9004 | 32.7818
Lu | 1.0701 | 0.94414 | 7.5575 | 6.6437 | 6.3976 | 6.5401 | 6.7232 | 5.7044 | 5.5183 | 36.597 | 35.4751 | 31.5829
Hf | 1.0371 | 0.91501 | 7.325 | 6.4125 | 6.1935 | 6.3145 | 6.4975 | 5.5021 | 5.3273 | 35.1732 | 34.0753 | 30.5329
Ta | 1.0057 | 0.88711 | 7.1029 | 6.1929 | 5.9989 | 6.0982 | 6.2809 | 5.3107 | 5.1445 | 33.8362 | 32.7632 | 29.4602
W | 0.97542 | 0.86045 | 6.8897 | 5.9822 | 5.8121 | 5.8926 | 6.0745 | 5.1267 | 4.9704 | 32.5971 | 31.5313 | 28.4324
Re | 0.94635 | 0.8348 | 6.6875 | 5.7802 | 5.635 | 5.6949 | 5.8765 | 4.9516 | 4.8032 | 31.3944 | 30.3567 | 27.4608
Os | 0.9184 | 0.81021 | 6.4924 | 5.5871 | 5.4631 | 5.5047 | 5.6865 | 4.7838 | 4.6433 | 30.2931 | 29.2517 | 26.5169
Ir | 0.89158 | 0.78656 | 6.3063 | 5.4034 | 5.3024 | 5.3244 | 5.5047 | 4.624 | 4.4912 | 29.2222 | 28.1694 | 25.6699
Pt | 0.86575 | 0.76385 | 6.1279 | 5.2263 | 5.1477 | 5.1518 | 5.3307 | 4.4707 | 4.3455 | 28.2106 | 27.2026 | 24.8219
Au | 0.84094 | 0.74195 | 5.957 | 5.0568 | 4.9983 | 4.9836 | 5.1638 | 4.3237 | 4.2052 | 27.2538 | 26.2403 | 24.0082
Hg | 0.81701 | 0.72097 | 5.7924 | 4.8938 | 4.8565 | 4.8237 | 5.0039 | 4.1836 | 4.0706 | 26.3598 | 25.3438 | 23.2556
Tl | 0.79399 | 0.70069 | 5.6344 | 4.7376 | 4.719 | 4.6699 | 4.8495 | 4.0484 | 3.9411 | 25.4777 | 24.4961 | 22.5048
Pb | 0.77178 | 0.68122 | 5.4833 | 4.587 | 4.5877 | 4.5228 | 4.7018 | 3.919 | 3.8169 | 24.6632 | 23.6839 | 21.8092
Bi | 0.75037 | 0.66244 | 5.3381 | 4.4426 | 4.4635 | 4.38 | 4.5591 | 3.7946 | 3.6981 | 23.8794 | 22.9057 | 21.1553
Po | 0.72972 | 0.64432 | 5.1981 | 4.3028 | 4.3416 | 4.2425 | 4.4215 | 3.675 | 3.584 | 23.1347 | 22.16 | 20.525
At | 0.70976 | 0.62687 | 5.0634 | 4.1698 | 4.224 | 4.1132 | 4.2916 | 3.5604 | 3.4744 | 22.4263 | 21.4693 | 19.9105
Rn | 0.69062 | 0.60989 | 4.9339 | 4.0416 | 4.1135 | 3.987 | 4.1361 | 3.4502 | 3.3692 | 21.76 | 20.8054 | 19.3253
Fr | 0.67201 | 0.59362 | 4.8092 | 3.9174 | 4.0066 | 3.8635 | 4.0442 | 3.3445 | 3.2682 | 21.109 | 20.1532 | 18.7552
Ra | 0.65401 | 0.57785 | 4.6888 | 3.7976 | 3.9029 | 3.7462 | 3.9235 | 3.2416 | 3.1673 | 20.4886 | 19.5406 | 18.2064
Ac | 0.63663 | 0.56257 | 4.5732 | 3.6823 | 3.8006 | 3.6319 | 3.8071 | 3.1432 | 3.0727 | 19.8899 | 18.9518 | 17.6725
Th | 0.61982 | 0.54787 | 4.4614 | 3.5712 | 3.7075 | 3.5225 | 3.6988 | 3.048 | 2.9815 | 19.3124 | 18.3916 | 17.1691
Pa | 0.60354 | 0.53363 | 4.3533 | 3.4642 | 3.6156 | 3.4176 | 3.5929 | 2.9569 | 2.8952 | 18.7735 | 17.858 | 16.6936
U | 0.58777 | 0.51985 | 4.2497 | 3.36 | 3.5265 | 3.3148 | 3.4908 | 2.869 | 2.8089 | 18.2466 | 17.3389 | 16.2391
Np |  |  | 4.1494 | 3.2597 | 3.4359 | 3.2164 | 3.3914 | 2.7837 | 2.7409 | 17.7866 | 16.854 | 15.7958
Pu |  |  | 4.0521 | 3.1628 | 3.3532 | 3.1206 | 3.2957 | 2.7016 | 2.6633 | 17.2974 | 16.3863 | 15.3678
Am |  |  | 3.9584 | 3.0692 | 3.2732 | 3.0284 | 3.2032 | 2.6222 | 2.5875
Cm
Bk
Cf
Es

## xraydata
X-ray Line energies in ev from Waldo program (reordered), Handbook of Chemistry and Physics, orig ref Bearden and Burr
Symbol | Ka | Kb | La1 | Lb1 | Lb2 | Lb3 | Lb4 | Lg1 | Lg2,3 | Ma | Mb | Mg
H | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
He | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Li | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Be | 100 | 100 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
B | 183 | 183 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C | 277 | 277 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
N | 392 | 392 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
O | 525 | 525 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
F | 677 | 677 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ne | 849 | 858 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Na | 1041 | 1071 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Mg | 1254 | 1302 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Al | 1487 | 1557 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Si | 1740 | 1836 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
P | 2014 | 2139 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
S | 2308 | 2464 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cl | 2622 | 2816 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ar | 2958 | 3191 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
K | 3314 | 3590 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ca | 3692 | 4013 | 341 | 341 | 341 | 341 | 341 | 341 | 341 | 0 | 0 | 0
Sc | 4091 | 4461 | 395 | 400 | 395 | 400 | 400 | 400 | 400 | 0 | 0 | 0
Ti | 4511 | 4932 | 452 | 458 | 452 | 458 | 458 | 458 | 458 | 0 | 0 | 0
V | 4952 | 5427 | 511 | 519 | 511 | 519 | 519 | 519 | 519 | 0 | 0 | 0
Cr | 5415 | 5947 | 573 | 583 | 573 | 585 | 585 | 583 | 585 | 0 | 0 | 0
Mn | 5899 | 6490 | 637 | 649 | 637 | 654 | 654 | 649 | 654 | 0 | 0 | 0
Fe | 6403 | 7058 | 705 | 719 | 705 | 721 | 721 | 719 | 721 | 0 | 0 | 0
Co | 6930 | 7649 | 776 | 791 | 776 | 792 | 792 | 791 | 792 | 0 | 0 | 0
Ni | 7478 | 8265 | 852 | 869 | 852 | 941 | 941 | 869 | 941 | 0 | 0 | 0
Cu | 8048 | 8905 | 930 | 950 | 930 | 1023 | 1023 | 950 | 1023 | 0 | 0 | 0
Zn | 8639 | 9572 | 1012 | 1035 | 1012 | 1107 | 1107 | 1035 | 1107 | 0 | 0 | 0
Ga | 9252 | 10264 | 1098 | 1125 | 1098 | 1197 | 1197 | 1125 | 1197 | 0 | 0 | 0
Ge | 9886 | 10982 | 1188 | 1219 | 1188 | 1294 | 1286 | 1219 | 1294 | 0 | 0 | 0
As | 10544 | 11726 | 1282 | 1317 | 1282 | 1388 | 1388 | 1317 | 1388 | 0 | 0 | 0
Se | 11222 | 12496 | 1379 | 1419 | 1379 | 1490 | 1490 | 1419 | 1490 | 0 | 0 | 0
Br | 11924 | 13291 | 1480 | 1526 | 1480 | 1596 | 1596 | 1526 | 1596 | 0 | 0 | 0
Kr | 12649 | 14112 | 1586 | 1637 | 1586 | 1697 | 1697 | 1637 | 1707 | 0 | 0 | 0
Rb | 13395 | 14961 | 1694 | 1752 | 1694 | 1827 | 1818 | 1752 | 2051 | 0 | 0 | 0
Sr | 14165 | 15836 | 1807 | 1872 | 1807 | 1947 | 1936 | 1872 | 2197 | 0 | 0 | 0
Y | 14958 | 16738 | 1923 | 1996 | 1923 | 2072 | 2060 | 1996 | 2347 | 0 | 0 | 0
Zr | 15775 | 17668 | 2042 | 2124 | 2042 | 2201 | 2187 | 2124 | 2503 | 0 | 0 | 0
Nb | 16615 | 18623 | 2166 | 2257 | 2367 | 2335 | 2319 | 2462 | 2664 | 0 | 0 | 0
Mo | 17479 | 19608 | 2293 | 2395 | 2518 | 2473 | 2456 | 2624 | 2831 | 0 | 0 | 0
Tc | 18367 | 20619 | 2424 | 2537 | 2630 | 2605 | 2592 | 2792 | 3000 | 0 | 0 | 0
Ru | 19279 | 21657 | 2559 | 2683 | 2836 | 2763 | 2741 | 2965 | 3181 | 0 | 0 | 0
Rh | 20216 | 22724 | 2697 | 2834 | 3001 | 2916 | 2891 | 3144 | 3364 | 0 | 0 | 0
Pd | 21177 | 23819 | 2839 | 2990 | 3172 | 3073 | 3045 | 3329 | 3553 | 0 | 0 | 0
Ag | 22163 | 24942 | 2984 | 3151 | 3348 | 3234 | 3203 | 3520 | 3746 | 0 | 0 | 0
Cd | 23174 | 26096 | 3134 | 3317 | 3528 | 3401 | 3367 | 3717 | 3951 | 0 | 0 | 0
In | 24210 | 27276 | 3287 | 3487 | 3714 | 3573 | 3535 | 3921 | 4161 | 0 | 0 | 0
Sn | 25271 | 28486 | 3444 | 3663 | 3905 | 3750 | 3708 | 4131 | 4377 | 0 | 0 | 0
Sb | 26359 | 29726 | 3605 | 3844 | 4101 | 3933 | 3886 | 4348 | 4600 | 0 | 0 | 0
Te | 27472 | 30996 | 3769 | 4030 | 4302 | 4120 | 4070 | 4571 | 4829 | 0 | 0 | 0
I | 28612 | 32295 | 3938 | 4221 | 4508 | 4313 | 4258 | 4801 | 5066 | 0 | 0 | 0
Xe | 29779 | 33624 | 4110 | 4417 | 4719 | 4513 | 4258 | 5037 | 5305 | 0 | 0 | 0
Cs | 30973 | 34987 | 4287 | 4620 | 4936 | 4717 | 4649 | 5280 | 5546 | 0 | 0 | 0
Ba | 32194 | 36378 | 4466 | 4828 | 5157 | 4927 | 4852 | 5531 | 5804 | 0 | 0 | 0
La | 33442 | 37801 | 4651 | 5042 | 5384 | 5143 | 5062 | 5789 | 6067 | 833 | 854 | 1027
Ce | 34720 | 39257 | 4840 | 5262 | 5613 | 5365 | 5277 | 6052 | 6332 | 883 | 902 | 1075
Pr | 36026 | 40748 | 5034 | 5489 | 5850 | 5592 | 5498 | 6322 | 6607 | 929 | 950 | 1127
Nd | 37361 | 42271 | 5230 | 5722 | 6089 | 5829 | 5722 | 6602 | 6892 | 978 | 997 | 1180
Pm | 38725 | 43826 | 5433 | 5961 | 6339 | 6071 | 5957 | 6892 | 7185 | 1029 | 1049 | 1237
Sm | 40118 | 45413 | 5636 | 6205 | 6587 | 6318 | 6196 | 7178 | 7496 | 1081 | 1100 | 1291
Eu | 41542 | 47038 | 5846 | 6456 | 6843 | 6571 | 6439 | 7480 | 7782 | 1131 | 1153 | 1346
Gd | 43000 | 48697 | 6057 | 6713 | 7103 | 6831 | 6687 | 7786 | 8096 | 1185 | 1209 | 1402
Tb | 44482 | 50229 | 6273 | 6978 | 7367 | 7096 | 6940 | 8102 | 8410 | 1240 | 1266 | 1461
Dy | 45998 | 51957 | 6495 | 7248 | 7636 | 7370 | 7204 | 8419 | 8733 | 1293 | 1325 | 1522
Ho | 47547 | 53877 | 6720 | 7525 | 7911 | 7652 | 7471 | 8747 | 9069 | 1348 | 1383 | 1576
Er | 49128 | 55681 | 6949 | 7811 | 8189 | 7939 | 7745 | 9089 | 9417 | 1406 | 1443 | 1643
Tm | 50742 | 57517 | 7180 | 8101 | 8468 | 8231 | 8026 | 9426 | 9754 | 1462 | 1503 | 1705
Yb | 52389 | 59370 | 7416 | 8402 | 8759 | 8537 | 8313 | 9870 | 10016 | 1521 | 1568 | 1765
Lu | 54070 | 61283 | 7656 | 8709 | 9044 | 8847 | 8606 | 10143 | 10485 | 1581 | 1631 | 1832
Hf | 55790 | 63234 | 7899 | 9023 | 9342 | 9163 | 8905 | 10516 | 10861 | 1645 | 1698 | 1895
Ta | 57532 | 65223 | 8146 | 9343 | 9645 | 9488 | 9212 | 10895 | 11247 | 1710 | 1766 | 1964
W | 59318 | 67244 | 8398 | 9672 | 9955 | 9819 | 9525 | 11286 | 11641 | 1775 | 1835 | 2035
Re | 61140 | 69310 | 8652 | 10010 | 10268 | 10160 | 9846 | 11685 | 12046 | 1843 | 1906 | 2107
Os | 63001 | 71413 | 8912 | 10356 | 10591 | 10511 | 10175 | 12095 | 12461 | 1910 | 1978 | 2182
Ir | 64896 | 73561 | 9175 | 10708 | 10912 | 10867 | 10511 | 12513 | 12883 | 1980 | 2054 | 2254
Pt | 66832 | 75748 | 9442 | 11071 | 11240 | 11231 | 10854 | 12942 | 13315 | 2051 | 2127 | 2331
Au | 68804 | 77984 | 9713 | 11442 | 11576 | 11610 | 11205 | 13382 | 13759 | 2123 | 2205 | 2410
Hg | 70819 | 80253 | 9989 | 11823 | 11914 | 11995 | 11563 | 13830 | 14214 | 2195 | 2283 | 2488
Tl | 72872 | 82576 | 10269 | 12213 | 12261 | 12390 | 11931 | 14292 | 14681 | 2271 | 2362 | 2571
Pb | 74969 | 84936 | 10552 | 12614 | 12612 | 12793 | 12306 | 14764 | 15159 | 2346 | 2443 | 2653
Bi | 77108 | 87343 | 10839 | 13024 | 12963 | 13210 | 12691 | 15248 | 15646 | 2423 | 2526 | 2735
Po | 79290 | 89800 | 11131 | 13447 | 13327 | 13638 | 13086 | 15744 | 16144 | 2501 | 2611 | 2819
At | 81520 | 92300 | 11427 | 13876 | 13698 | 14067 | 13482 | 16251 | 16653 | 2580 | 2695 | 2906
Rn | 83780 | 94870 | 11727 | 14316 | 14066 | 14512 | 13989 | 16770 | 17173 | 2659 | 2781 | 2994
Fr | 86100 | 97470 | 12031 | 14770 | 14441 | 14976 | 14307 | 17300 | 17704 | 2741 | 2871 | 3085
Ra | 88470 | 100130 | 12340 | 15236 | 14825 | 15445 | 14747 | 17849 | 18268 | 2824 | 2961 | 3178
Ac | 90884 | 102850 | 12652 | 15713 | 15224 | 15931 | 15198 | 18408 | 18830 | 2909 | 3053 | 3274
Th | 93350 | 105609 | 12969 | 16202 | 15606 | 16426 | 15643 | 18983 | 19406 | 2996 | 3146 | 3370
Pa | 95868 | 108427 | 13291 | 16702 | 16003 | 16930 | 16104 | 19568 | 19985 | 3082 | 3240 | 3466
U | 98439 | 111300 | 13615 | 17220 | 16407 | 17455 | 16575 | 20167 | 20599 | 3171 | 3337 | 3563
Np | 0 | 0 | 13944 | 17750 | 16840 | 17989 | 17061 | 20785 | 21110 | 3253 | 3433 | 3663
Pu | 0 | 0 | 14279 | 18294 | 17255 | 18541 | 17556 | 21417 | 21725 | 3345 | 3531 | 3765
Am | 0 | 0 | 14617 | 18852 | 17677 | 19106 | 18063 | 22065 | 22361 | 0 | 0 | 0
Cm | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bk | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cf | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Es | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Edge energies in ev from Waldo fortran code edge.for, from Handbook of Chemistry and Physics, original source Bearden and Burr
Symbol | K | L1 | L2 | L3 | M1 | M2 | M3 | M4 | M5 | N1 | N2 | N3
H | 14 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
He | 25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Li | 55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Be | 111 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
B | 188 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C | 284 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
N | 402 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
O | 532 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
F | 685 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ne | 867 | 45 | 18 | 18 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Na | 1072 | 63 | 31 | 31 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Mg | 1305 | 89 | 51 | 51 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Al | 1560 | 117 | 73 | 73 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Si | 1839 | 149 | 99 | 99 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
P | 2146 | 189 | 132 | 132 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
S | 2472 | 229 | 165 | 165 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cl | 2822 | 270 | 202 | 200 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ar | 3203 | 320 | 247 | 245 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
K | 3607 | 377 | 296 | 294 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ca | 4038 | 438 | 350 | 346 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Sc | 4493 | 500 | 407 | 402 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ti | 4966 | 564 | 462 | 456 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
V | 5465 | 628 | 521 | 513 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cr | 5989 | 695 | 584 | 575 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Mn | 6539 | 769 | 651 | 640 | 84 | 49 | 49 | 3 | 3 | 0 | 0 | 0
Fe | 7112 | 846 | 721 | 708 | 93 | 54 | 54 | 4 | 4 | 0 | 0 | 0
Co | 7709 | 926 | 794 | 779 | 101 | 54 | 54 | 3 | 3 | 0 | 0 | 0
Ni | 8333 | 1008 | 872 | 855 | 112 | 68 | 68 | 4 | 4 | 0 | 0 | 0
Cu | 8979 | 1096 | 951 | 931 | 120 | 74 | 74 | 2 | 2 | 0 | 0 | 0
Zn | 9659 | 1194 | 1043 | 1020 | 136 | 87 | 87 | 8 | 8 | 0 | 0 | 0
Ga | 10367 | 1298 | 1142 | 1115 | 158 | 107 | 103 | 17 | 17 | 0 | 0 | 0
Ge | 11103 | 1414 | 1248 | 1217 | 180 | 128 | 121 | 29 | 29 | 0 | 0 | 0
As | 11867 | 1527 | 1359 | 1323 | 204 | 146 | 141 | 41 | 41 | 3 | 3 | 3
Se | 12658 | 1654 | 1476 | 1433 | 232 | 168 | 162 | 57 | 57 | 6 | 6 | 6
Br | 13474 | 1782 | 1596 | 1550 | 257 | 189 | 182 | 70 | 69 | 27 | 5 | 5
Kr | 14326 | 1921 | 1727 | 1675 | 289 | 223 | 214 | 89 | 89 | 24 | 11 | 11
Rb | 15200 | 2065 | 1864 | 1804 | 322 | 247 | 239 | 112 | 110 | 29 | 15 | 14
Sr | 16105 | 2216 | 2007 | 1940 | 358 | 280 | 269 | 135 | 133 | 38 | 20 | 20
Y | 17038 | 2373 | 2156 | 2080 | 394 | 312 | 300 | 160 | 157 | 45 | 26 | 26
Zr | 17998 | 2532 | 2307 | 2222 | 430 | 344 | 331 | 182 | 180 | 51 | 29 | 29
Nb | 18986 | 2698 | 2465 | 2371 | 468 | 378 | 363 | 207 | 205 | 58 | 34 | 34
Mo | 20000 | 2866 | 2625 | 2520 | 505 | 410 | 392 | 230 | 227 | 62 | 35 | 35
Tc | 21044 | 3043 | 2793 | 2677 | 544 | 445 | 425 | 256 | 253 | 68 | 39 | 39
Ru | 22117 | 3224 | 2967 | 2838 | 585 | 483 | 461 | 284 | 279 | 75 | 43 | 43
Rh | 23220 | 3412 | 3146 | 3004 | 627 | 521 | 496 | 312 | 307 | 81 | 48 | 48
Pd | 24350 | 3604 | 3330 | 3173 | 670 | 559 | 532 | 340 | 335 | 86 | 51 | 51
Ag | 25514 | 3806 | 3524 | 3351 | 718 | 602 | 571 | 373 | 367 | 95 | 63 | 56
Cd | 26711 | 4018 | 3727 | 3538 | 770 | 651 | 617 | 411 | 404 | 108 | 67 | 67
In | 27940 | 4238 | 3938 | 3730 | 826 | 702 | 664 | 451 | 443 | 122 | 77 | 77
Sn | 29200 | 4465 | 4156 | 3929 | 884 | 756 | 714 | 493 | 485 | 137 | 89 | 89
Sb | 30491 | 4698 | 4380 | 4132 | 944 | 812 | 766 | 537 | 528 | 152 | 98 | 98
Te | 31814 | 4939 | 4612 | 4341 | 1006 | 870 | 819 | 583 | 572 | 168 | 110 | 110
I | 33169 | 5188 | 4852 | 4557 | 1072 | 931 | 875 | 631 | 619 | 186 | 123 | 123
Xe | 34561 | 5453 | 5104 | 4782 | 1143 | 999 | 937 | 685 | 672 | 208 | 147 | 147
Cs | 35985 | 5714 | 5359 | 5012 | 1217 | 1065 | 998 | 740 | 726 | 231 | 172 | 162
Ba | 37441 | 5989 | 5624 | 5247 | 1293 | 1137 | 1062 | 796 | 781 | 253 | 192 | 180
La | 38925 | 6266 | 5891 | 5483 | 1361 | 1204 | 1123 | 849 | 832 | 270 | 206 | 191
Ce | 40443 | 6549 | 6164 | 5723 | 1435 | 1273 | 1185 | 901 | 883 | 290 | 223 | 207
Pr | 41991 | 6835 | 6440 | 5964 | 1511 | 1337 | 1242 | 951 | 931 | 305 | 236 | 218
Nd | 43569 | 7126 | 6722 | 6208 | 1575 | 1403 | 1297 | 1000 | 978 | 315 | 243 | 224
Pm | 45184 | 7428 | 7013 | 6459 | 1650 | 1471 | 1357 | 1052 | 1027 | 330 | 242 | 242
Sm | 46834 | 7737 | 7312 | 6716 | 1723 | 1541 | 1420 | 1106 | 1080 | 346 | 266 | 247
Eu | 48519 | 8052 | 7617 | 6977 | 1800 | 1614 | 1481 | 1161 | 1131 | 360 | 284 | 257
Gd | 50239 | 8376 | 7930 | 7243 | 1881 | 1688 | 1544 | 1217 | 1185 | 376 | 289 | 271
Tb | 51996 | 8708 | 8252 | 7514 | 1968 | 1768 | 1611 | 1275 | 1241 | 398 | 310 | 285
Dy | 53789 | 9046 | 8581 | 7790 | 2047 | 1842 | 1676 | 1333 | 1295 | 416 | 332 | 293
Ho | 55618 | 9394 | 8918 | 8071 | 2128 | 1923 | 1741 | 1392 | 1351 | 436 | 344 | 307
Er | 57486 | 9751 | 9264 | 8358 | 2207 | 2006 | 1812 | 1453 | 1409 | 449 | 366 | 320
Tm | 59390 | 10116 | 9617 | 8648 | 2307 | 2090 | 1885 | 1515 | 1468 | 472 | 386 | 337
Yb | 61332 | 10486 | 9978 | 8944 | 2398 | 2173 | 1950 | 1576 | 1528 | 487 | 397 | 344
Lu | 63314 | 10870 | 10349 | 9244 | 2491 | 2264 | 2024 | 1639 | 1589 | 506 | 410 | 359
Hf | 65351 | 11271 | 10739 | 9561 | 2601 | 2365 | 2108 | 1716 | 1662 | 538 | 437 | 380
Ta | 67416 | 11682 | 11136 | 9881 | 2708 | 2469 | 2194 | 1793 | 1735 | 567 | 465 | 405
W | 69525 | 12100 | 11544 | 10207 | 2820 | 2575 | 2281 | 1872 | 1809 | 595 | 492 | 425
Re | 71676 | 12527 | 11959 | 10535 | 2932 | 2682 | 2367 | 1949 | 1883 | 625 | 518 | 444
Os | 73871 | 12968 | 12385 | 10871 | 3049 | 2792 | 2457 | 2031 | 1960 | 654 | 547 | 468
Ir | 76111 | 13419 | 12824 | 11215 | 3174 | 2909 | 2551 | 2116 | 2040 | 690 | 577 | 494
Pt | 78395 | 13880 | 13273 | 11564 | 3296 | 3027 | 2645 | 2202 | 2122 | 722 | 609 | 519
Au | 80725 | 14353 | 13734 | 11919 | 3425 | 3148 | 2743 | 2291 | 2206 | 759 | 644 | 545
Hg | 83102 | 14839 | 14209 | 12284 | 3562 | 3279 | 2847 | 2385 | 2295 | 800 | 677 | 571
Tl | 85530 | 15347 | 14698 | 12657 | 3704 | 3416 | 2957 | 2485 | 2389 | 846 | 721 | 609
Pb | 88005 | 15861 | 15200 | 13035 | 3851 | 3554 | 3066 | 2586 | 2484 | 894 | 764 | 645
Bi | 90526 | 16388 | 15711 | 13419 | 3999 | 3696 | 3177 | 2688 | 2580 | 938 | 805 | 679
Po | 93105 | 16939 | 16244 | 13814 | 4149 | 3854 | 3302 | 2798 | 2683 | 995 | 851 | 705
At | 65730 | 17493 | 16785 | 14214 | 4317 | 4008 | 3426 | 2909 | 2787 | 1042 | 886 | 740
Rn | 98404 | 18049 | 17337 | 14619 | 4482 | 4159 | 3538 | 3022 | 2892 | 1097 | 929 | 768
Fr | 101137 | 18639 | 17907 | 15031 | 4652 | 4327 | 3663 | 3136 | 3000 | 1153 | 980 | 810
Ra | 103922 | 19237 | 18484 | 15444 | 4822 | 4490 | 3792 | 3248 | 3105 | 1208 | 1058 | 879
Ac | 106755 | 19840 | 19083 | 15871 | 5002 | 4656 | 3909 | 3370 | 3219 | 1269 | 1080 | 890
Th | 109651 | 20472 | 19693 | 16300 | 5182 | 4830 | 4046 | 3491 | 3332 | 1330 | 1168 | 967
Pa | 112601 | 21105 | 20314 | 16733 | 5367 | 5001 | 4174 | 3611 | 3442 | 1387 | 1224 | 1007
U | 115606 | 21757 | 20948 | 17166 | 5548 | 5182 | 4303 | 3728 | 3552 | 1441 | 1273 | 1045
Np | 118678 | 22427 | 21601 | 17610 | 5723 | 5366 | 4435 | 3850 | 3666 | 1501 | 1328 | 1087
Pu | 121818 | 23097 | 22266 | 18057 | 5933 | 5541 | 4557 | 3973 | 3778 | 1559 | 1372 | 1115
Am | 125027 | 23773 | 22944 | 18504 | 6121 | 5710 | 4667 | 4092 | 3887 | 1617 | 1412 | 1136
Cm | 128220 | 24460 | 23779 | 18930 | 6288 | 5895 | 4797 | 4227 | 3971 | 1643 | 1440 | 1154
Bk | 131590 | 25275 | 24385 | 19452 | 6556 | 6147 | 4977 | 4366 | 4132 | 1755 | 1554 | 1235
Cf | 135960 | 26110 | 25250 | 19930 | 6754 | 6359 | 5109 | 4497 | 4253 | 1799 | 1616 | 1279
Es | 139490 | 26900 | 26020 | 20410 | 6977 | 6574 | 5252 | 4630 | 4374 | 1868 | 1680 | 1321